MULTI-ETHNIC STUDY OF ATHEROSCLEROSIS

EXAM 3

Field Center Procedures

Manual of Operations

March 17, 2004

Multi-Ethnic Study of Atherosclerosis

Exam 3 ~ Field Center Procedures ~ Manual of Operations

Table of Contents

SECTION 1:  OVERVIEW AND CLINIC EQUIPMENT	5

SECTION 2:  OVERVIEW OF EXAM 3 COMPONENTS	11

2.1	Summary of Exam 3 components	11

2.1.1	Exam 3 Procedures:	11

2.1.2	Exam 3 Forms	11

2.1.3	Surveillance Phone Call Forms and Questionnaires:	11

2.1.4	Events Forms	13

SECTION 3:  CLINIC QUESTIONNAIRES & EXAMINATION	14

3.1  	Pre-Exam Activities	14

3.1.1 	Preparatory Activities Prior to Participant Arrival	14

3.1.2	Instruction to Participants Before the Clinic Visit	15

3.2  	Examination Guidelines	16

3.3  	Clinic Reception & Clinic Check Off Forms	18

3.3.1	Clinic Reception	18

3.3.2	Clinic Check Off Sheet	20

3.4  	Interviews - Questionnaires	21

3.4.1 	Interviewing Guidelines and Techniques	21

3.4.2	Participant Tracking Form	23

3.4.3  	Personal History	26

3.4.4  	Medical History	31

3.4.5 	Medications	35

3.4.6  	Physical Activity	41

3.4.7  	Health and Life	47

3.4.8  	Family History	51

3.4.9	Sleep History	58

3.4.10	Eye History	61

3.4.11	Neighborhood Activities	66

3.4.12  	Residential History	76

3.5  	Clinic Examinations	80

3.5.1 	Anthropometry	80

3.5.2  	Seated Blood Pressure	83

3.5.3  Ankle/Arm Blood Pressure	87

3.5.4  	Phlebotomy & Spot Urine Collection (see Laboratory, section 3.7)	90

3.6	MESA-Eye	91

3.7  	Laboratory	92

3.8   Carotid Ultrasound	108

3.8.1 General Guidelines	108

3.8.2 Right Common Carotid Dynamic Acquisition	111

3.8.3  Carotid IMT	121

3.9  	Carotid MRI – non-contrast and Gadolinium-Enhanced	140

3.10 	 Computed Tomography (CT)	141

3.11   CT Calcium Scoring  and Diameter of the Aorta	142

SECTION 4:  ALERTS	148

SECTION 5:  REPORTING PARTICIPANTS’ RESULTS	150

SECTION 6:  DATA MANAGEMENT	151

SECTION 7:	EVENTS SURVEILLANCE	155

7.1 	 Follow-up Call	155

7.2 	 Continuing Participant Surveillance	155

7.1 	 Use of Proxy	156

APPENDIX 1:  Handwriting Examples for Teleform	157

SECTION 1:  OVERVIEW AND CLINIC EQUIPMENT

1.1	MESA Exam 3 Brief Overview

Welcome to Exam 3 of the Multi-Ethnic Study of Atherosclerosis (MESA)! You all have done a wonderful job recruiting for and completing the crucial first two parts (Exams 1 and 2) of this vitally important national research project that will influence diagnosis and treatment of cardiovascular diseases over the next several decades.

Now onward! As we continue on, keep the above in mind when you perform the procedures on those gracious and willing participants. The ultimate value of this study depends on you and on the quality of the data that you collect.

1.1.1	Purpose and Objectives

1.1.1.1	The purpose of MESA is to study subclinical cardiovascular disease (CVD).  Subclinical diseases are those detected by non-invasive procedures, such as ECG and ultrasound, before they have produced any clinical signs or symptoms.

1.1.1.2	The primary objectives of MESA are to determine:

	What factors influence the progression of mild subclinical disease to more severe subclinical disease?

	What factors influence the progression of subclinical disease to clinical disease?

1.1.1.3	The secondary objectives of MESA are to:

	Assess ethnic, age, and gender differences in subclinical CVD prevalence and risk of progression

	Describe the interrelationships of established risk factors for subclinical CVD with new risk factors that are identified by the MESA study

	Develop population-based methods, suitable for application in future screening and intervention studies, to identify asymptomatic individuals who are at high risk for subclinical CVD.

1.1.2	Brief Description of the MESA study

1.1.2.1	MESA has been designed as a large and complex long-term study.  MESA has recruited a cohort of nearly 6,800 men and women ages 45–85. Approximately 38% of the cohort are Caucasian, 28% African-American, 23% Hispanic, and 11% Asian, predominantly of Chinese descent.

1.1.2.2	The cohort was recruited from the following six Field Centers (ID number in parentheses):

	Wake Forest University, Winston-Salem, NC (3)

	Columbia University, New York, NY (4)

	Johns Hopkins University, Baltimore, MD (5)

	University of Minnesota, Minneapolis-St. Paul, MN (6)

	Northwestern and Loyola Universities, Chicago, IL (7)

	UCLA, Los Angeles, CA (8)

1.1.2.3	In Exam 1, there were 6 Reading and Labs Centers.  Exam 3 will have the addition of an Eye center.  These centers are located at:

	Harbor-UCLA (CT RC)

	Johns Hopkins University (MRI RC)

	Tufts-New England Medical Center (Boston, Ultrasound RC)

	Wake Forest University (ECG RC)

	University of Minnesota (Lipid Lab)

	University of Vermont (Central Blood Lab)

- University of Vermont (Lab center)

- University of Wisconsin (Eye Center)

1.1.2.4	The Coordinating Center at the University of Washington (Seattle) continues to coordinate all aspects of the study, including development of the Manual of Operations, development of data entry and management software, quality control, and statistical analysis of the data for publication in academic journals.

1.1.2.5	Funding for MESA is provided by the National Heart, Lung, and Blood Institute (NHLBI), a division of the National Institutes of Health. The Project Office from NHLBI also provides scientific leadership to the study.

1.1.2.6	The entire MESA project will take 10 years to complete. The first 18 months of the study were devoted to protocol development, staff training, and pilot testing. MESA Exam 1 (i.e. the “Baseline Exam”) took place over a two-year period, beginning in late July 2000. The second and third exams will each require about 18 months to complete, and the fourth exam will require two years to complete. During these four exams, participants will also be contacted by telephone every 6–9 months to determine whether any medical events have occurred. The final 18 months of the study will be dedicated to close out and data analysis for publication.

1.1.3	Events

In MESA, an “event” is the development of a medical condition requiring participant hospitalization or other specified types of treatment, or the death of a participant.  In MESA, we are particularly interested in collecting data about myocardial infarction (MI or “heart attack”), stroke, transient ischemic attack (TIA or “mini-stroke”), angina, congestive heart failure (CHF), peripheral vascular disease (PVD), and death. When an event occurs, we will collect a separate set of data based on hospital and physician records and on interviews with participants or their proxies.

1.1.4	How to use this manual

This manual contains step-by-step instructions for completing all of the components in Exam 3 of the MESA study.  Many of the steps and clinic procedures are the same as for “MESA Exam 1 and 2”; thus where appropriate the “MESA Manual of Operations ~ Field Center Procedures ~ Exam 1” and “MESA Manual of Operations ~ Field Center Procedures ~ Exam 2”should be consulted.

You should carefully study all sections that relate to procedures that you will be performing, and you should keep this manual as well as the exam 1 manual of operation handy as references. If you have questions about anything in the Manual, please direct them to your Study Coordinator.

1.2	  Supplies and Equipment

Most supplies and equipment needed for Exam 3 are similar to that of Exams 1 and 2.  New equipment required is listed first, followed by those the FCs should already be familiar with.  Each clinic should be equipped with the following supplies and equipment:

    - MESA- Retina (Retinal Photography)

- Canon CR6-45NM fundus camera with a digital Canon D-60 camera back

- Laptop computer, with EyeQSL software, provided by Digital Healthcare Inc.

- Motorized instrument table

- Pneumatically adjustable stools for both photographer and subjects.

- Lens cotton – One 4 oz. Box (long-grain Red Cross sterile cotton batting)

- Lens cleaning alcohol – One 8oz. bottle (100% alcohol)

- Facial tissues – Six boxes (200 tissues/box)

- Spare view lamp (#BH3-3277) – One

- Canon camera fuses (125V, 4 amp) – Two

- 100 CDs and jewel cases, CD labels

- Shipping envelopes and preprinted shipping labels, provided by the Ocular Epidemiology Reading Center (OERC).

- Instrument table and stools.

1.2.2	MESA – Vision (Refraction)

- Nidek Auto Lensmeter (model LM-990A)

- Nidek Autorefractor/Keratometer (model ARK-760A)

- 3 foot Cable from lensmeter to autorefractor (LM-03-0808)

- 10 foot Cable from autorefractor to computer (TRS-10-0809)

- Computer

- Surge Protector

- Spectacles of known prescription (NEI will provide to each clinic)

- Penlight

- Blower brush

- QC folder

- Mirror (for use in inserting & removing contacts)

- Plastic tray (to place contact lens case in while removing or inserting contacts)

- Unisol 4 saline solution – NOTE:  You must write the date opened on the bottle.  The bottle should be discarded 30 days after opening.

- Clarity lens cleaner

- Contact lens cases

- Alcohol wipes

- Chin rest tissue

- ARK printer paper

- Lensmeter printer paper

- Lint-free Kimwipes

- Absorbond lens wipers

- Cotton tipped applicators

- Surgical gloves

QC items (to be circulated among the various clinics) include:

Test eyes/steel balls (ARK9-00-ARKJ-4)

1.2.3		Equipment Already Provided by the Coordinating Center

	2 ECG machines and modems

	Computers, scanners, and printers

1.2.4	Phlebotomy and Laboratory Supplies being Provided by the Central Laboratory

	5 mL SCAT-1 tubes

	8 mL Cell Preparation tubes

	Cryogenic vials (0.5 mL, 2.0 mL, 10.0 mL)

	Nalgene Freezing container #2 'Mr. Frosty' (for cell prep) (Nalge#5100-001)

	15 mL centrifuge tubes (for cell prep)

	Freezing Media A and B

	PBS (phosphate buffered saline)

	Acetic acid (for urine)

	ACD/dextran (for red cell membranes)

- Urine receptacle package

1.2.5	Equipment & Supplies Already Purchased and available at Field Centers

1.2.5.1	Examination Equipment

	Three Dinamap Blood Pressure Machines, Model PRO 100

	Rapid Cuff Inflation E20

	AG 101 Airsource with BP 10D & SC 12 D and PAK5 Positioning Kit

	Detecto Platform Balance, # 1915-339W Scale in lbs/kg

	Stadiometer (Accu-Hite Measure Device w/level Bubble)

	Two Gulick II anthropometric tapes in cm, Model 67021

	Wall mounted mirror for verifying tape measurement

	Wall thermometer

	Doppler machine, # EN 50 Nicolet LE 100 5 megahertz vascular probe

	Wall mounted mercury sphygmomanometer (Baum)

	Blood pressure cuffs (pediatric, adult, large, and thigh)

1.2.5.2	Laboratory and Phlebotomy Equipment

	Two Centrifuges with temperature control, 2,000 g-force minimum, swinging bucket, and test tube holders (adaptors), e.g., IEC Centra CL3R with CL3R swinging bucket rotor (#243) and aerocarrier tube holder for IEC 243; adapters need are #6561E, 6561E, 6562E, 6566E.

	Harvard Trip Balance / Pan balance, VWR # 12344-051

	Water tubes for balancing the centrifuge, VWR # 21008-102

	Freezer (-70 C or colder)

	Refrigerator for storage of special blood tubes, media, etc (can be a household fridge, should not be used to store food, from Sears or a similar store)

	Test tube racks / cryovial racks (Simport # T315)

	Fixed volume pipettes with tips (MLA) and regular adjustable pipettes (Rainin, Finn, etc) with tips.  Volumes needed to pipette: 225 ul, 0.5 ml, 1.0 ml (200 to 1000 ul), 3 to 5 ml, 9 ml, and 14 ml.

	Graduated cylinder, Lab Safety 25 ml Nalgene 9A-22763

	Blood tube rocker (Thermolyne Labquake Tube Shakers C400-110)

	Blood tube racks, Fisher Scientific # 60914764

	Stopwatches or timers, VWR # 62344-756

1.2.5.3	Other Equipment

	Refrigerator for storing food (i.e., a household fridge, should not be used to store blood samples or chemicals, can be ordered from Sears or a similar store)

	Ice machine, Ice Machine Sales Service # 351-FAV-161

	Sharps containers (wall mounted and table-top), sharps disposal Lab safety # 9A-25703

	Biohazard waste containers – Lab Safety Supply # 6 gallon container - # 9A-29190

	Emergency eye wash station – Lab Safety Supply # 9A-2121S

	Personal protective equipment for phlebotomy and processing: face shield, Lab safety # 9A-1402, splash shields, other Personal Protective Equipment as needed.

1.2.5.4	Examination Supplies

	Disposable ECG electrodes (Unitrace blue heart electrode)

	Single use thermometers

	Ultrasound transmission gel,

	Gloves (latex, vinyl, or others, any vendor)

	Kim wipes

	Pregnancy test kit

	Peel-off china marker

1.2.5.5	Phlebotomy and Laboratory Supplies

	Butterfly needles (21 G) with luer adapter

	Vacutainer barrels

	Tourniquets #CK1126

	Alcohol prep pads # 10-3001

	Gauze (2x2) Kendall # 1806

	Surgical tape - paper tape 3 M durapore

	Band-aids (first –aid) # 1290033

	Blood collection tubes:

- 10 ml Serum (red-top) tubes

- 10 ml EDTA (purple-top) tubes

- 4.5 ml Citrate (blue-top) tubes

	Polypropylene (or polystyrene) disposable test tubes (10-15 mL) for pooling samples

	Revco Boxes and dividers (10 x 10 and 7 x 7 grids) (Revco box # 5954 with Revco 10x10 grid # 5958, Revco box # 5956 with Revco 7x7 grid# 5959)

	Isopropyl alcohol

	Distilled water

	10% bleach solution (or approved biohazard disinfectant)

1.2.5.6	Shipping Supplies

	Styrofoam/insulated boxes (Polyfoam Packer # 355-CS for frozen shipments, #33- 12KD for refrigerated shipments)

	Ziplock bags for freezer boxes

	Dry ice

	Ice (gel) packs

	Absorbent material (old newspaper or paper towels) for layering between dry ice

	Elastic bands for freezer boxes

	Packing tape

1.2.5	Vendors

Vendors for non MESA-Retina & MESA-Vision materials are the same as for Exam 1.  Please refer to MESA Manual of Operations ~ Field Center Procedures ~ Exam 1.

Vendors for MESA-Retina materials are as follows:

CANON USA, Inc.

Irvine, California or Itasca, Illinois office

contacts are:

Mark Scheckle, 15955 Alton Parkway,

Irvine California, 92618-3616

telephone number: 949-753-4193

e-mail: mscheckel@cusa.canon.com

Tom Penkala, Canon USA, Inc., 100 Park Blvd.,

Itasca, Illinois, 60143-2693

telephone number: 630-250-6230

e-mail: tpenkala@cusa.canon.com

The photography consultants, Michael Neider and Hugh Wabers and Katie Nigl can be reached at 608-263-9858.

Vendors for MESA-Vision materials are as follows:

- All NIDEK equipment as well as paper & supplies for the equipment is supplied by:

- Marco Technologies

- Mr. Bill Lavelle is in charge of all government contract work.

- 11825 Central Parkway Jacksonville FL 32224,

- telephone number: 1-800-874-5274).

- The other non-consumables (penlight, mirror, plastic bin) and the consumables can be obtained anywhere - grocery stores, local mall optician, or pharmacy.

SECTION 2:  OVERVIEW OF EXAM 3 COMPONENTS

  - Summary of Exam 3 components

2.1.1	Exam 3 Procedures:

  - Anthropometry

  - Seated Blood Pressure

  - Ankle-Arm Blood Pressure

  - Phlebotomy

  - Spot Urine

  - Ultrasound IMT

  - Carotid MRI, with and without contrast

  - CT Examination

  - MESA –Vision, Refraction

  - MESA – Retina, Retinal photography

2.1.2	Exam 3 Forms

2.1.2.1	Clinic Examination Data Forms and Questionnaire

- Clinic Reception

- Clinic Check-off Sheet

- Questionnaire Data Forms

- Anthropometry

- Seated Blood Pressure

- Ankle-Arm Blood Pressure

- Personal History

- Medical History

- Medications History

- Neighborhood Activities

- Physical Activity History

- Health & Life

- Residential History

- Residential History Supplemental Page

- Eye History

      - Reading Center Completion Forms

              - MESA – Vision Completion, Refraction

              - MESA – Retina Completion, Retinal photography

              - Phlebotomy Completion

              - Lab Processing

              - Urine Collection Completion

              - MRI Exam, exclusion & completion

              - CT Exam Completion

              - Ultrasound IMT

2.1.3	Surveillance Phone Call Forms and Questionnaires:

- Participant tracking

- Contact Log

- Contact Cover Sheet

- General Health

- General Health – Death

- Death Information

- Specific Medical Conditions

- Specific Medical Procedures

- Other Admissions

2.1.4	Events Forms

  - Initial Notification of Potential Event/Death

  - Events Eligibility

  - Events Eligibility Addendum

  - Hospital Abstraction: Cardiac/PVD

  - Hospital Abstraction: Stroke/TIA

  - Physician Questionnaire: Cardiac/PVD

  - Physician Questionnaire: Stroke/TIA

  - Physician Questionnaire: Cardiovascular Death

  - Events Contact Log

  - Cardiac/PVD Interview

  - Interview for Stroke/TIA Symptoms

  - Informant Interview

  - Narrative for Stroke

  - Cardiac/PVD Review Form

  - Mortality Review Form

  - Stroke/TIA Review Form

SECTION 3:  CLINIC QUESTIONNAIRES & EXAMINATION

3.1  	Pre-Exam Activities

3.1.1 	Preparatory Activities Prior to Participant Arrival

For participants due to have an Exam 3 visit, s/he should be contacted at minimum one week prior to their planned appointment date.  During this contact, the entire set of Pre-Exam 3 Phone Call questionnaires will be completed (see section 7), an Exam 3 appointment scheduled, and if possible, appointments for CT and/or MRI.

The purpose of completing the Pre-Exam 3 Phone Call questionnaire is to obtain the most recent information on the participants’ health status as well as events and recent radiation (CT, X-ray) exposures, which may affect their ability to participate in the Exam 3 CT component.

Preparatory steps prior to clinic visits are the same as for MESA Exam 1 and 2.  In addition, the Participant Data Report (a summary of data from the participant’s Exam 1 and 2 visits) should be printed, the Pre-Exam 3 Phone Call forms should be available, and both should be reviewed for completeness prior to the participant’s arrival.  Any information missing should be noted and obtained.

1.	Calendars, Report, and Forms

1.1	The data manager should print out the Daily Calendar showing clinic visits scheduled for the following day. The Calendar lists the preferred language of each person scheduled for a clinic visit, which will help you to determine how many sets of language-specific forms you will need for the day.

1.2	For each participant on the Daily Calendar, print the Participant Data Report and have his/her Surveillance Phone Call forms available.  Some questionnaires and procedure will require information on the Participant’s Data Report.

1.3	Print several sets of the forms that will be completed for interviews and procedures. These will be pre-printed with participant IDs, which will be matched to an individual when s/he begins the exam. Forms can be printed as far in advance as you would like, and it is advisable to have enough printed forms on hand for at least several days of clinic visits. Forms will be printed in English, Chinese, or Spanish. For each participant, gather all the forms required for a visit, including the informed consent, medical release, and screening (telephone interview) form, and place into the participants binder. Make sure his/her ID matches the binder.

2.	Supplies and Equipment

2.1	Set up vacutainer and aliquoting tubes on the racks and attach the pre-printed labels to the tubes. Place all phlebotomy supplies on the blood drawing table (21 g, luer adaptor, vacutainer barrel, tourniquet, alcohol pad, gauze 2x2, surgical tape, Band-Aid).

2.2	Make sure that the examination rooms are clean and have clean linen.

2.3	Prepare the participants’ gowns or (scrubs) and slippers.

2.4	Prepare the examination room for seated BP, anthropometry measurement, etc. Check all instruments that will be used for the examination.

2.5	Make sure breakfast will be available for the participants.

3.	Staffing

3.1	Prepare staff assignment sheet and make sure everyone knows his/her responsibilities. This is particularly important if you schedule a large number of participants on a given day.

  - If participants have been scheduled for CT or MRI, make sure that the respective technicians have the participant schedule and forms.

3.1.2	Instruction to Participants Before the Clinic Visit

Mail instructions to the participant 7–10 days before the clinic visit and explain them over the telephone when you schedule the visit. If possible, make a reminder call to the participant the day before the clinic visit and reiterate the instructions. (If the participant is acutely ill—e.g. “flu” or bronchitis—when you make this reminder call, tell him/her not to come to the clinic. Arrange to contact him/her again to reschedule when he/she has recovered.) Before the examination, make sure the participants understand the following instructions.

1.	Participants must fast (except water) for at least 12 hours before the examination. Instruct them to consume dinner at least 12 hours before their scheduled appointment at the clinic. Only water and prescription medications are permitted from dinner until the start of the examination the next morning.

2.	Participants should avoid heavy exercise during the 12 hours before the visit.

3.	Participants should not smoke on the morning of the visit.

4.	Participants should bring all current medications, both prescription and over-the-counter, including vitamin preparations, dietary supplements, and herbal remedies to the clinic. If the participant forgets to bring the medications, schedule another clinic visit to obtain this information or collect the information when the participant returns for imaging procedures.

- 5.	Participants should bring the name and complete address of their personal physician or health plan, particularly if they wish to have examination results sent to that provider.

3.2  	Examination Guidelines

I.	MESA EXAM 3 GUIDELINES

The MESA Exam 3 will be scheduled over a 16-month period, beginning March 4, 2004. The examination will include several questionnaires and procedures (i.e. anthropometry, blood pressure measurement, fasting blood collections, Eye exams, etc.). Selected participants will have a CT exam and/or an MRI exam. We estimate that the complete examination including CT and/or MRI will require about six hours. The examination may be performed in one day or over several days. However, every effort should be made to perform all the components of the examination within a four-week period. Clinics may vary the exam sequence if needed; but all components of Exam 3 must be completed.

- A pregnancy test should be performed within 48 hours before the CT exam in all women of child-bearing potential who are scheduled to have a CT exam. This could be done in the clinic or in the Radiology Department.

2.	Questionnaires and clinic procedures should be performed before the CT and/or MRI exams.

3.	Anthropometry and blood collection should be performed while the participant is fasting. (If participant is not fasting, record date and time he/she last ate or drank.)Blood pressure measurement should be done before venipuncture. CT, MRI, Eye exams, and questionnaires do not require fasting.

4. 	Blood drawing should be done after a 12-hour fast and before 12:00 noon.

II.	EXAMINATION ORDER

Guidelines for clinic order are similar to that of Exam 1 and 2.  Many elements are left to the discretion of the individual field center.  A few key points are reiterated here:

1.	Resting blood pressures should be obtained after the subject has been in the seated position for at least five minutes.

2.	Venipuncture should be performed in the fasting state after blood pressure measurement. If a participant comes to the clinic non-fasting, perform exam components that do not require fasting, and schedule the participant for another clinic visit for fasting blood collection.

3.	Questionnaires may be administered at any time during the examination. During the interviews, make every effort to avoid distractions, ensure privacy, and maintain confidentiality for the participant. Do not conduct interviews during the snack or in the waiting area in the clinic.

III.	GUIDELINES FOR EXAMINATION OF DIABETIC PARTICIPANTS

1.	Diet-controlled diabetics must fast overnight and are treated the same as non-diabetics.

2.	Diabetics taking oral hypoglycemic medications or insulin must fast overnight (unless a bedtime snack was prescribed by their physician) and to come to the clinic without taking their hypoglycemic medication. They should bring their morning medication dose with them to the clinic. Schedule all known diabetics taking oral hypoglycemic medications or insulin for examination as early as possible (before 9 a.m.). Draw fasting blood samples promptly on arrival at the clinic (after measuring blood pressure). Immediately following venipuncture, serve breakfast and instruct participants to take hypoglycemic medication as prescribed.

3.3  	Clinic Reception & Clinic Check Off Forms

3.3.1	Clinic Reception

I. 	Purpose

As with Exam 1, the Exam 3 Clinic Reception Form is the means by which important information collected during the reception process is recorded for scanning into the clinic database. This form must be completed accurately to allow for the tracking of the participant’s progress through the study and to provide a link to the sampling and recruitment data.

II.	Methods

1. 	General Instructions

1.1	The clinic reception process is very important in setting the participant’s frame of mind for the rest of the exam day. Greet each participant warmly as soon as he/she arrives at the clinic. (If a participant arrives at the clinic acutely ill—e.g., “flu” or bronchitis—do not continue with the clinic examination. Make arrangements to contact him/her to reschedule the appointment after he/she has recovered.)

1.2	Instruct the participant to read the informed consent documents carefully, answer the questions at the end, and sign it (see section on informed consent). Once the participant has completed and signed the inform ed consent, the clinic visit will begin.

1.3	Ask the participant if he/she has any questions. After you have answered any questions, give the participant a gown (or robe) and slippers and take him/her to a dressing room to change. Provide a locker or other safe place for the participant’s clothing and any other items that need to be stored. Complete the Clinic Reception form. This concludes the Clinic Reception phase of the clinic visit.

2.	Specific Instructions for Completing the Clinic Reception Form

2.1	Because the Clinic Reception Form serves several important functions, a second staff member must review it before scanning.

2.2	The Participant ID, Acrostic, Birth Date, QC ID, and Language spoken are preprinted on the form in the upper right corner.

2.3	In the Visit Date fields, record a two-digit month and day and a four-digit year (e.g. 09/25/2004). Whenever possible, the second appointment date for Exam 3, if any, should be entered on the form before scanning. If the second appointment cannot be scheduled before the form is scanned, the data manager should update the database as soon as the information is available.

2.4	In the Local Medial/Hospital Identification Number field, record the participant’s local hospital or medical record number. This number is not required, however.

2.5	In the Informed Consent Responses area, transcribe the responses from the participant’s consent form by filling the appropriate bubble.

  - In the Reception Interview area, ask the participant when they last ate or drank.  Record the information in military (24:00 hours) time.  Record the time, again in military time, when this form is being completed.

- The 2 times given are less than 8 hours apart; reschedule the visit or the fasting component of the exam.

- Then ask if the participant has been ill in the last seven days (e.g. cold, flu, fever, vomiting.)  If the participant responds “Yes”, inform the participant that the clinic exam cannot be completed at this time, thank him/her and reschedule another visit.

  - In the Interviewer ID field, record the ID number of the person who administered the Clinic Reception Form.

  - In the Reviewer ID field, record the ID number of the person who reviewed the Clinic Reception Form.

3.3.2	Clinic Check Off Sheet

I. 	Purpose

As with Exams 1 and 2, the Exam 3 Clinic Check Off Sheet is the means to ensure that all parts of Exam 3 are completed or scheduled. This form must be completed accurately to allow for the tracking of the participant’s progress through Exam 3.

II.	Methods

1. 	General Instructions

1.1	The form should be filled out as completely as possible during the exam visit day.  Make sure to schedule dates for the participant’s CT, MRI, and/or Ultrasound IMT examinations, if appropriate.

1.2	Provide the participant with appointment reminders for the CT, MRI, and/or Ultrasound IMT exams, if appropriate.

1.3	Ask the participant if he/she has any questions at exam exit time. After you have answered any questions, thank the participant. This concludes the Clinic check off sheet and the clinic visit day for the participant.

2.		Specific Instructions for Completing the Clinic Check Off Sheet

  - The process in completing this form is the same as Exams 1 and 2.

  - For each step/procedure, make certain to fill out the start and end time in the appropriate fields. Note any comments and record the Tech ID.

  - Visit day progress order should be as recommended for the first five steps/procedures.  The remainder of the procedures could be performed in any order except the MESA eye component.  The Eye History questionnaire should always precede the eye examinations. MESA Vision (refraction) should be completed before MESA Retina (retinal photograph).

  - Check and note those participants selected to have CT and/or MRI in Exam 3. Call and schedule CT and/or MRI examinations if have not done so and prepare appointment reminders to give to the participant on exit.  Record the date and time for the appointment/s in the appropriate field.

  - At the end of the form, ask the participant for his/her “preference in method of future MESA contacts.”  Choices include: Telephone call, Letter, E-mail. Fill in the box that applies.  If the participant prefers e-mail, ask for the e-mail address and transcribe on the line provided.

  - Do Not Scan this form!

3.4  	Interviews - Questionnaires

3.4.1 	Interviewing Guidelines and Techniques

I.	GENERAL INTERVIEW INFORMATION

1.	Interviewer bias is any preference or inclination that creates a systematic difference between responses obtained by different interviewers. It can be affected by:

	Respondent's perception of the interviewer and his/her reaction to that

	Interviewer's perception of the respondent and his/her reaction to that

2.	Characteristics of a good interview

2.1	The interviewer creates a friendly, but businesslike atmosphere.

2.2	The respondent is at ease. Keep these factors in mind:

 	the respondent may view a female interviewer as less threatening.

	the respondent may view a much older interviewer as judgmental.

2.3	The interviewer obtains the answer to the question that is asked by:

	proper use of probes and repeating a question rather than interpreting it.

2.4	The interviewer obtains clarification of confusing answers.

2.5	The interviewer gives only neutral responses to the respondent's answers.

2.6	The interviewer accurately records responses.

3.	Specific skills required for interviewers

3.1	The ability to ask questions at the correct pace and in a conversational tone.

3.2	A thorough knowledge of the questions and response categories (this will keep the interview flowing smoothly).

3.3	Knowledge of how and when to use probes.

3.4	The ability to think as an interviewer and to temporarily put aside other roles (e.g., researcher, health care provider).

3.5	The ability to maintain a positive attitude about the interview (this lets the respondent know that the interview is important).

3.6	The ability to keep some level of control over the interview process (e.g., by rewarding the respondent for answering questions but not for other behavior).

II.	INTERVIEWING TECHNIQUES

1.	Standardized Interviewing Technique

1.1	MESA is a collaborative study being conducted through six field centers located throughout the United States.  In order to produce data that can be considered collaborative, MESA study designers must develop and use standardized approaches to train interviewers and collect information about participants. Standardization is achieved by using scripts in training, training supervisors centrally, establishing qualifications for supervisors, reviewing collected data, taping and reviewing interviews, and, finally, observing interviewers in the field.

  - It is critically important that interviewers read the sections in this manual that is applicable to the questionnaire they will be administering.

III.	The INTERVIew

The following procedures are recommended for a successful interview:

1.	Prior to the visit prepare all materials (e.g., appropriate report, forms, identification, stamped-self-addressed envelopes) that will be necessary for the interview.

2.	Find an area where both you and the participant can talk and write comfortably with minimal distractions.

3.	Make sure that the participant understands the questions and that you are interpreting the responses accurately. Do this by restating what you think the participant is telling you. At the same time, be careful not to impose your interpretations on the interview questions or the participant's comments.

4.	Convey your interest in the participant’s thoughts and feelings, but do your best to keep him/her focused on the interview questions. When the participant strays from a question, try to use what he/she is saying to redirect the conversation back to the interview questions. Give positive reinforcement for direct answers. If necessary, set time limits at the outset of the interview to encourage the participant to stay on track.

5.	Participants may try to convince you to answer certain questions for them.  Let the participants know that you are interested in their answers.

6.	Be aware of any hearing and vision impairments and their effects on the participant's understanding of the interview questions. If necessary, read the interview questions to participants who have visual impairments or limited reading ability.

7.	Encourage, but do not force, participants to answer to all questions.

8.	If non-participants are present during the visit, address the participant directly and do not encourage conversation with other parties. If necessary, ask that you and the participant be left alone for a brief time to complete the questionnaire.

9.	Be able to adapt to interruptions. Let the participant know that you are willing to continue the interview after the interruptions are completed.

10.	Make the interview a positive experience for the participant. React favorably to answers and give compliments, when appropriate.

11.	Give the participant clear information about when the next clinic visit will be conducted and follow through with the plans that you make.

3.4.2	Participant Tracking Form

I.	Purpose

The Tracking Form is similar to and is an update of Exam 1 information collected.  It allows us to collect information (name, address, telephone number, and email address) on the participant, his/her health care provider(s), and any proxies or contacts he/she may designate. We will use this information to contact and communicate with the participant and his/her physician(s), proxies, or other contacts.  In Exam 3, this information needs to be reviewed and updated with any changes since form completion during Exam 2.

II.	Materials/Equipment

The participant’s personal address book, a phone book, and a computer will all be useful in helping participants find and record the information asked for on the Tracking Form.

III.	Definitions

1.	Proxy. A person designated by the participant to knowledgeably answer questions about the participant, in the event that he/she is unable to answer. A proxy may, but does not have to, live with the participant and should be familiar with the status of the participant’s health.

2.	Contact. A person designated by the participant who may be relied upon to know the participant’s whereabouts. A contact does not live with the participant, but always knows how to get in touch with him/her.

3.	Email address. A computer address where electronic mail is received e.g. bozo@clownmail.com.

4.	Second surname. Another last name used by the participant. Some participants (e.g., some members of the Hispanic population) use two last names. Also, some married women use both their maiden name and their husband’s last name.

IV.	Methods

General instructions

  - Exam 2 data for this questionnaire will be pre-printed on the tracking form for the interviewer to compare and update any changes per participant response.

  - Current information is essential for maintaining contact with participants and for communicating with their proxies, contacts, and health care providers. You should emphasize to the participants that the tracking form needs to be updated as completely and accurately as possible. Also encourage participants to designate as proxies only those people who are familiar with the status of their health, because it is the proxies who will answer health-related questions if the participant is unable to.

  - Participants should verify and clearly print, in ink and in capital letters, all changes in the appropriate spaces. The participant should not use a nickname in place of a full, legal name. He/she should provide an area code with each phone number, even if within the local calling area. Boxes/spaces for items of information that are not applicable should be left blank. You should verify health care provider information using a local telephone directory. Obtain missing information over the phone.

Specific Instructions:

Ask the participant as outlined below to obtain participant information:

Section A.  Participant Information:

Begin with, “Please VERIFY your name, address, telephone number(s), and email address (if you have one).” Note any changes in section “A”, “Changes” column. Then, ask the participant the following questions:

A1. “Do you plan to change your name within the next year?” (Note that this is question as well as questions A2 & A3 are not written in the form)

If “no”, continue to question A2.

If “yes”, ask, “what will your new last name be?”

Record the information in section “A”, “Changes” column.  Then, continue to the next question, A2.

A2. “Do you plan to be out of this area for an extended period of time (a month or longer) within the next year?”

If “no”, continue to question A3.

If “yes”, ask, “approximately when will you leave and when will you return?”

Record the month/year for both in section “A”, “Changes” column. Then, continue to the next question, A3.

A3. “Will there be a change in your local address within the next three months?”

If “no”, continue to section B, contacts/proxies.

If “yes”, ask, “what will your new address be?”

Record the street address, city, state, and ZIP code in section “A”, “changes” column. Then, continue with section B.

Section B.  Contacts/Proxies Information:

“Please update the following information on people who are familiar with the status of your health AND who could help us contact you, if necessary. If possible, please include one person who lives with you and one who does not.”

The participant should provide as much information as possible. Assist him/her, if necessary, in obtaining information.  Record any changes regarding ‘proxies’ in the lines to the right of that particular ‘proxy’.

If a ‘proxy’ was used to obtain the information on this form, the interviewer should check the appropriate box under the name of the ‘proxy’ person.

If the ‘proxy’ used to obtain the information on this form is not listed among the Exam 1 proxies, ask the ‘proxy’ for his/her name, relationship to the participant, complete address, and telephone number.  Fill the information in the spaces provided at the end of section B.

Section C.  Health Care Providers Information:

“Please update the following information about your health care providers i.e. a clinic, doctor, nurse, or physician’s assistant who provides your usual medical care?”

If the participant does not have a health care provider then the form is complete.  Thank the participant.

If the participant has a health care provider, record any changes in the “Changes” column.  The form is now complete.  Thank the participant.

Remember to fill in the interviewer ID, reviewer ID and data entry ID at the end of the page.

3.4.3  	Personal History

I.	Purpose

The Personal History questionnaire is used to collect information on socio-economic status (SES) and smoking and drinking habits, all of which are related to an individual’s risk of cardiovascular disease.

II.	Methods

General instructions:

This is a self-administered questionnaire. Provide the participant with the form and a pencil and give brief instructions for completion. If the participant is unable to self-administer the questionnaire, then a MESA staff member will administer the questionnaire.

Ask the participant to try to answer all questions, unless instructed to skip the question. Remind him/her to request assistance from a staff member if anything is unclear. Most participants should be able to complete the questionnaire on their own. However, if the participant expresses or appears to have difficulty reading or comprehending the questions, offer to help and make arrangements for an interviewer administered version in the appropriate language.

Specific instructions:

- Instruct the participant to read the questions and their instructions carefully then fill out all questions, except those he/she is instructed to skip as a result of his/her response to a specific question.

- If he/she is unsure about an exact answer (e.g., for “average number of drinks per week”), tell him/her to give a best estimate.

- In questions where the participant is asked about number of times use per day, instruct him/her to fill in “00” if use is less than once a day.

Participant Information (questions 1–10)

Responses to Questions 1-10 are used to update information previously provided during the Exam 1 of the MESA study.  The participant will begin the questionnaire with the introduction below:

This form is intended to collect information about your background and lifestyle which may impact your risk of cardiovascular disease. Please complete all items except those which you are specifically instructed to skip. If you are unsure about the answer to a specific question, please estimate the answer to the best of your ability. If you have a question about a particular item, please write a small ‘x’ in the margin of the form, making sure not to write it near any of the response bubbles, and then ask a staff member for clarification of those items after you have completed the rest of the form.

1a.	Has your employment status changed since your MESA clinic visit on (Exam 2 visit date)?  Choose Yes or No.

If no, skip to question #6.

If yes, continue with 1b.

1b.	Choose one of the following that best describes your current occupation. Select a choice and fill in the appropriate bubble.

If homemaker, not working outside the home: Did you previously work outside the home? Choose Yes or No.

If no, skip to question #6.

If yes, continue with the next question, Q#2.

Questions 2–5 are used to compute industry and occupation codes. The actual coding will be done at a later time. The interviewer may or may not be the person who does the coding, so it is important that job information is complete, accurate, and legible. (Note:  Questions 2–5 will not be scanned.  Only the final codes will be included in the database.)  This set of questions begins with:

“In Questions #2-5, please refer to your current or most recent occupation.

2.	For whom do/did you work? (name of company, business, organization, or other employer.) If you are not working now, please respond regarding your main occupation before you stopped working. Write answer in the space provided.

3.	What type of business or industry is/was this? (e.g. hospital, newspaper publishing, mail order house, auto repair shop, bank, etc.) Write answer in the space provided.

4.	What kind of work do/did you do or what is/was your job title? (e.g. registered nurse, personnel manager, auto mechanic, accountant, grinder operator, etc.) Write answer in the space provided.

5.	What are/were your most important activities or duties? (e.g., patient care, directing hiring policies, repairing automobiles, reviewing financial records, operates grinding mill, etc.) Write answer in the space provided.

Questions #6-11 ask about the participant’s family finances.  It begins with:

The following questions have to do with family finances. We know from other research that financial strain is common and very important to consider in understanding people’s health. The following questions will be used to help give us a picture of the various financial situations experienced by persons participating in the MESA study. Any information you provide is strictly confidential and will be used for research purposes only.

6.	Below is a list of income groups. Please tell me which group best represents your total combined family income for the past 12 months. This includes the total income, before taxes, earned in the past year by all family members living with you. Please include money from jobs; net income from business, farm, or rent; pensions, dividends, welfare, social security payments, and any other money received by you or any other family member living with you. Select a choice and fill in the appropriate bubble.

7a.	Including yourself, how many people are supported by the income listed in the previous question? Write a number in the box provided.

7b.	Including yourself, how many of these are:

	Children under 18?

	Adults 65 and over?

Write a number in each box. Enter 00 if no one in that age category is supported by the given income.

8.	This question is about the house or apartment where you live. Do you: Choose one of the choices given and fill in the appropriate bubble.

(“Rent” includes renting a room in an owner-occupied dwelling or renting a room in a rooming house. “Other living arrangements” include, for example, living for free with parents or friends or receiving free rent in exchange for apartment management, yard work, etc.)

- Questions 9-13 are intended to measure socioeconomic characteristics, which may be related to cardiovascular health.

- Questions 9- 11 obtain additional information on the participant's financial status. Participants may have questions about why we want to know this. Research has shown that income alone is often not enough to characterize a person's financial status, and that things about financial status may be very important to a person's health. Obtaining this information will be helpful to the study in understanding the causes of cardiovascular disease. You may emphasize that this information will be strictly confidential (no personal identifiers are on the form). However, if the participant remains uncomfortable about answering the questions he or she should feel free to skip them.

- 9. 	Do you or your family have investments such as stocks, bonds, mutual funds, retirement investments, or other investments?

- Choices are “Yes,” or “No”.  Select the best response and fill in the bubble.

- 10. 	Do you or your family own any land, business property, apartments, or houses other than the one in which you now live?

- Choices are “Yes, own,” “Currently buying,” or “No.” Select the best response.

- 11. 	Do you or your family own a car?

- Choices are “Yes, 1 car,” “Yes, more than 1 car,” or “No.” Select the best response.

Questions 12 & 13 ask about medical care.

- Where do you usually go for medical care?

- Choices are “Doctor’s office or clinic,” “Hospital Emergency room,” or “Others.”  Choose the best response and fill in the appropriate bubble.

- If “other,” specify in the box provided. Mark “other” only if the response clearly does not fit one of the given responses. For example, an urgent care clinic would be included in the “doctor’s office or clinic” category.

13.	To help pay for your medical care, do you now have...? Choose all that apply and fill in the appropriate bubble(s). If “other,” specify in the box provided. Mark “other” only if the response clearly does not fit one of the given responses.

Smoking or tobacco product usage, Questions 14–26

The goal of these questions is to identify non-smokers, former smokers, and current smokers of any tobacco product and to document participants’ smoking habits. These questions were developed from NHANES III, the National Health Interview Survey, and ARIC. If the participant feels uncomfortable with these questions, please reassure him/her that all collected data is strictly confidential.  This section begins with the following introductory script:

The following questions are about your use of tobacco and alcohol. They will help us better understand the role of smoking and alcohol use in the risk of cardiovascular disease.

14.	Which of the following best describes your current smoking status?

Choose the appropriate response and fill in the bubble.

If never smoked, skip to question 17 and continue with the questionnaire.

15.	Have you smoked cigarettes during the last 30 days?

Choose “yes” or “no” and fill in the bubble.

If “no”, skip to question #19.  If “yes” continue with #16

16.	On average, about how many cigarettes a day do you smoke? Provide the number of cigarettes smoked per day then skip to question 18.

The participant should record 00 if the average number of cigarettes per day is less than one. Make sure participants record the information in number of cigarettes per day.  If a participant answers in number of packs per day, recalculate into number of cigarettes per day (1 pack = 20 cigarettes).

17.	Current non-smokers only: During the past year about how many hours per week were you in close contact with people when they were smoking? (e.g. in your home, in a car, at work, other close quarters etc.)

Provide number of hours per week.

This question applies to current non-smokers and former users of any kind of tobacco product.  The goal of the question is to obtain information on passive exposure to cigarette smoke (excluding cigars, pipes, etc.) in any type of close quarters during the past 12 months.  Record the number of hours in a typical week; do not include isolated or atypical situations, such as holiday gatherings or short-term house guests who smoke. If participants do not remember the exact amount of time, ask them to give their best estimate. Record 00 if participant was exposed to less than 1 hour of per week.

18.	Since your last MESA visit, have you used any other tobacco products (e.g. cigars, pipes, snuff, chewing tobacco)? The goal of this question is to obtain information about the use of other tobacco products. All participants should answer this question.

Choices are “yes” or “no”. If no, skip to question #27.

19.	Have you smoked cigars during the last 30 days?

Choices are “yes” or “no”. If no, skip to question #21.

20.	On average, about how many cigars a day do you smoke?

Provide number of cigars per day. Record 00 if the average number of cigars per day is less than one.

21.	Have you smoked a pipe during the last 30 days?

Choices are “yes” or “no”. If no, skip to question #23.

22.	On average, about how many pipefuls a day do you smoke?

Provide number of pipefuls per day. Record 00 if the average number of pipefuls per day is less than one.

23.	Have you chewed tobacco, such as Redman, Levi Garret, Beechnut, during the last 30 days?

Choices are “yes” or “no”. If no, skip to question #25.

24.	On average, about how many times a day do you chew tobacco?

Provide the number of times per day. Record 00 if the average number of uses per day is less than one.

25.	Have you used snuff, such as Skoal, Skoal Bandits, or Copenhagen during the last 30 days?

Choices are “yes” or “no”. If no, skip to question #27.

26.	On average, about how many times a day do/did you use snuff?

Provide number of times per day.  Record 00 if the average number of uses per day is less than one.

Alcohol Use, Questions 27–32

The goal of these questions is to update information previously provided on alcohol use.  Changes in alcohol consumption may be important factors in understanding atherosclerosis formation and diseases caused by atherosclerosis.  Frequency of alcohol intake is determined by usual weekly intake. The serving sizes are different for beer (12 oz glass, can, or bottle), wine (3.5 oz), and hard liquor (1.5 oz = 1 “shot”). Mixed drinks or cocktails (e.g. martini, margarita) are prepared with hard liquor and other beverages, such as soda, seltzer, syrup, etc. Questions are in boldface type; explanations, if needed, follow in standard type; possible answers follow in parentheses.

27.	Do you presently drink alcoholic beverages?

Choices are “yes” or “no”. If no, you are now finished with this questionnaire.

28. 	How many glasses of red wine do you usually have per week?

(1 serving = 3.5 oz glass, 1 bottle = 750ml = 8 glasses)

Provide the average number of drinks per week. Record 00 if the participant has had less than one drink per week.

29.	How many glasses of white wine do you usually have per week?

(1 serving = 3.5 oz glass, 1 bottle = 750ml = 8 glasses)

Provide the average number of drinks per week.

Record 00 if the participant has had less than one drink per week.

30.	How many cans, bottles, or glasses of beer do you usually have per week?

(1 serving = 12 oz glass, 1 bottle = 355ml = 1 glass)

Provide the average number of 12 oz drinks per week.

Record 00 if the participant has had less than one drink per week.

31.	How many drinks of liquor or mixed drinks do you usually have per week?   (1 serving of liquor = 1.5 oz shot-glass, or one mixed drink)

Provide the average number of drinks per week.

Record 00 if the participant has had less than one drink per week.

32.	In the past month what is the largest number of drinks you had in one day?  Provide the largest number of drinks in one day. The largest number of drinks includes all the wine, beer, and hard liquor consumed within a single 24-hour period in the past month.

At this point, the participant has completed the questionnaire. MESA staff will review the questionnaire for completeness, clarify any question that were not answer then complete the questionnaire by filling out the box “For MESA Field Center Use Only:”

- If the form was self-administered, check for completeness.

- Mark if form was self-administered or interviewer-administered.

- Record Interviewer or Reviewer ID.

- Record Data Entry ID.

3.4.4  	Medical History

I.	Purpose

The Medical History identifies the participant’s medical conditions and provides other information that may:

	be used to adjust for co-morbidity;

	characterize the participant's access to medical care; and

	characterize family history of CVD.

II.	Methods

General instructions:

This is an interviewer-administered questionnaire. Questions should be read to the participant verbatim as they appear on the form to ensure standardization.  In addition, any introductory and transitional wording should be read verbatim.

For most questions, possible responses are “Yes”, “No,” ‘Don’t Know,” and/or ‘Not Applicable” or “N/A” (not applicable). A few other questions have choices as indicated. Have the participant choose the appropriate response/s for each question.

Do not probe to make interpretations about a participant’s specific symptoms. Ask questions as written and record answers as given by the participant.

Specific instructions:

Begin the questionnaire by reading to the participant the following introduction:

The following are some questions about your medical History.  Questions refer to things that happened since your last MESA visit on ___________.  Please answer to the best of your knowledge.

1.	How would you say your health currently compares with other persons of your age?  Choices include: “better,” “same,” or “worse.” Choose the appropriate response to the best estimate.

Questions 2-7 pertain to conditions the participant has been told he or she has by a doctor since the last MESA visit. The participant should choose “Yes” or “No” if he/she is fairly sure about the diagnosis and “Don't Know” if he/she believes he/she might have been told about the diagnosis but is not sure.  If the person is cared for primarily by a health care provider other than a physician, such as a nurse practitioner, try to determine that the diagnosis was made in a medical setting and, if so, include the response.

For any question that asks about the age at diagnosis or when treatment was begun, assist the participant in converting the year to age, if necessary.

Has a doctor told you that you have developed any of the following since your last MESA visit on [-----]?

2.	Emphysema?  This includes "chronic bronchitis," "chronic obstructive pulmonary disease," or "COPD." Bubble in the appropriate response.

3. 	Asthma? Bubble in the appropriate response.

4. 	Arthritis? This includes “osteoarthritis,” “wear-and-tear arthritis,” “rheumatoid arthritis,” and forms of arthritis associated with other diseases. Bubble in the appropriate response.

5. 	Rheumatic heart disease or heart valve problems? Bubble in the appropriate response.

6. 	Liver disease?

If no, go to the next question.

If yes, “which type?”  Choices are: “Cirrhosis” or “Hepatitis”.

a. 	Cirrhosis (can also be called “liver failure”) is a chronic condition (that can develop and progress over many years) and may be associated with jaundice (yellow eyes and skin), weight loss, abdominal and leg swelling, and skin changes.

b. 	Hepatitis is an acute condition (occurs shortly after the infection) and usually is accompanied by jaundice and sometimes with abdominal pain, which may be the only signs of infection. Verify that Hepatitis was diagnosed by a health professional.

If yes to “Hepatitis”, the select the type(s) of Hepatitis:  Select all that apply.

	Hepatitis A (sometimes called "infectious hepatitis")

	Hepatitis B (sometimes called "serum hepatitis")

	Hepatitis C, D and E (may also be transmitted by blood products)

- Other

- Don’t know

7. 	Kidney disease? Bubble in the appropriate response.

Questions 8-9 pertain to how the participant feels about him/herself when compared to others of his/her own age.  The participant should be encouraged to estimate and answer Yes or No. The participant may choose “Don't Know” if he/she cannot give a yes or no answer.

8.	When walking on level ground, do you get more breathless than people your own age?

9.	When walking up hills or stairs, do you get more breathless than people your own age?

10. 	Do you ever have to stop walking because of breathlessness?

11. 	Do you ever get pain in either leg or buttock while walking?

If no, go to question #12 and continue with the questionnaire.

If yes, answer the following:

a. 	Does this pain ever begin when you are standing still or sitting?

b. 	In what part of your leg or buttock do you feel the pain? Choices include “pain includes calf/calves” or “pain does not include calf/calves.”

c.	Do you get it if you walk uphill or hurry?

d.	Do you get it if you walk at an ordinary pace on the level?

e.	Does the pain ever disappear while you are walking?

f.	What do you do if you get it when you are walking? Choices include: “stop or slow down” or “continue on.”

g.	What happens to the pain if you stand still? Choices include: “relieved” or “not relieved.”

If relieved, how soon? Select “10 minutes or less” or “more than 10 minutes.”

h.	Is this pain predominantly in the right side, left side, or in both legs? Select one of the choices.

12.	Since your last MESA clinic visit, have you had swelling (“edema”) of your feet or ankles? For women, exclude edema during pregnancy. Note: when a swollen extremity is pressed with a finger, an imprint or pit remains temporarily.

If no, continue with question #13.

If yes, ask the following: “Did it tend to come on during the day and go down overnight?”

13.	Since your last MESA clinic visit, have you had to sleep on two or more pillows to help you breathe? Some people may have to sleep in a chair to help them breath better; count this instance as “yes.”

14. 	In the past two weeks, have you had any of the following: This question is asked to determine if the participant has had some type of inflammatory condition. Choices for response to each of the following include: “Yes,” “No,” and “Don’t Know.”

a.	Fever

b. 	Cold, flu, or sore throat

c.	Urinary infection (also called “bladder infection”)

d.	Seasonal allergy, such as hay-fever

e.	Bronchitis

f.	 Sinus infection or sinusitis

g. 	Pneumonia

h.	Gums bleeding while brushing or flossing (include “periodontal disease” and “gingivitis”)

i.	Tooth infection requiring antibiotics and/or root canal

j.	Flare-up of gout

k.	Flare-up of arthritis

15.	Approximately how many times have you been treated with antibiotics in the last 12 months? If you don’t remember the exact number, please give your best estimate.

If none, record “00”.  If unable to give a number, then choose “Don’t Know”.

If the participant has been on daily antibiotics for an extended period, record “99” and record a comment in the space provided.

16.	Are you taking aspirin on a regular basis? Examples of "regular" are daily, every other day, and weekly.  If the participant says less than once a week, record “no”.

If yes, how many days a week are you taking aspirin? Give number of days/week.

At this point, Men are done with the questionnaire.

Reproductive History—for women only—

It is important to first check the participant’s printed clinic report and see if the answers to the following questions have already been	obtained and on record.  If the answer is available, skip as indicated below.

- Have you taken birth control pills since your last MESA clinic visit? (Y/N/DK)

If “no,” or “don’t know”, bubble in the response and continue with question #18.

If yes, ask the following: “Please estimate the total number of months that you took birth control pills since your last MESA clinic visit?”  Bubble in the response and continue with question #18.

18.	Have you had a hysterectomy (surgery to remove your uterus/womb)? (Y/N/DK)

- If the answer is available on the participant’s summary report, check the box above the question and skip to the next question, #19.

If “no,” or “don’t know”, bubble in the response and continue with question #19.

If yes, continue with question 18a:

18a.	“At what age?”  Record the response and continue with question #19.

19. 	Have you had surgery to remove your ovaries? Removal of the ovaries might have been in conjunction with a hysterectomy.

If the answer is available on the participant’s summary report, check the box above the question and skip to the next question, #20.

If “no,” or “don’t know”, bubble in the response and continue with question #20.

If yes, continue with question 19a & 19b:

19a. “At what age?” Record the response in the boxes.

19b. “How many ovaries were removed?” Select 1 or 2.  Continue with question #20.

20. 	Have you had a menstrual period in the last 12 months?

If the answer is available on the participant’s summary report, check the box above the question and skip to the next question, #21.

If “no,” or “don’t know”, bubble in the response and continue with question #21.

If yes, continue with question 20a:

20a.	“How many periods have you had in the last 12 months?” Record the response in the boxes and continue with question #21.

21.	Since your last MESA visit, have you taken hormone replacement therapy?

If no, questionnaire completed.

If yes, ask both 21a and 21b:

21a.	“Are you currently using hormone replacement therapy?”

If yes, “at what age did you begin?”

If no, at what ages did you take hormones? Provide age started and stopped.

21b.	“Which type of therapy were you on?” (Common estrogen-only preparations are Premarin or Estratab; common estrogen+progestin regimen is Premarin plus Provera, Estratab plus Provera, Prempro, or Premphase.)

Select “estrogen-only,” or “estrogen+progestin,” or “other types of hormone replacement”

After completion of the form, a technician should check to make sure all questions were answered and attempt to complete by asking the participant about specific skipped questions. Then complete the questionnaire by filling in the box “for MESA Field Center Use Only.”

3.4.5 	Medications

I.	Background and Rationale

1.	The MESA Exam 2 Medications Questionnaire is almost identical to that in the “MESA Exam 1”. The Medications Form is designed to enable collection of data on participants’ use of all types of medications, both prescription and non-prescription, including supplements. Information about participants’ use of medications is collected at the initial (baseline) clinic visit and at follow-up visits. The participant is asked to bring to the clinic containers for all medications used during the two weeks prior to the visit. The interviewer then transcribes the name of each medication, its strength, and for prescription medications, frequency of administration from the containers onto the data collection form.  As the information is transcribed, the interviewer queries the participant about actual usage of each medication.

2.	Collecting this information will allow us to describe medication use and any changing patterns of use over time, and may help us ascertain the effect of medications on the progression of atherosclerosis in this study population. It will be important to know what medications each participant is taking, in order to assess and perhaps attempt to explain subsequent participant events and any change in the degree of disease detected at follow-up visits.

II.	Materials and Equipment

Current version of the Physician’s Desk Reference (PDR)

III.	Definitions

1.	Time frame: All prescription and over-the-counter medications and supplements used during the two weeks prior to the clinic visit should be included.

2.	Prescription medication: Medication for which a prescription was written by a physician or physician assistant and dispensed by a pharmacist or a physician.

3.	Non-prescription or over-the-counter medication: Medication or supplements purchased without a prescription.

4.	It should be noted that occasionally a physician would write a prescription for a non-prescription medication. In that case, the medication should be recorded as prescription. If, however, the physician recommends a medication, rather than actually writing a prescription for it, it should be recorded as non-prescription.

IV.	Methods

This is an interviewer-administered questionnaire.  Questions should be read to the participant verbatim as they appear on the form to ensure standardization.  In addition, any introductory and transitional wording should be read verbatim.

1.	Obtaining medication containers. A letter is sent to the participant before the clinic visit that includes instructions regarding medication containers. The participant is asked to bring to the clinic containers for all supplements, prescription, non-prescription medications and herbal medicines taken during the two weeks prior to the clinic visit.

2.	Medication use interview. Prior to beginning the interview, place all medications in front of the participant. When asking the participant about a particular medication, show the container to the participant, keeping the other medications in view.  Always conclude the interview by asking the participant if any other medications have been taken during the previous two weeks. If the participant remembers other medications, record the name, strength and frequency administered for each one in as much detail as possible. If you are unsure about the accuracy of the participant’s responses, schedule a telephone interview to verify the prescription label information. At the end of the visit, make sure to return all medications and other personal belongings to the participant. Guidelines for completing the Medications Form follow:

Section A.	Medication Reception

As you know, the Multi-Ethnic Study of Atherosclerosis will be describing all medication its participants are using, both prescription and over-the-counter. These include pills, liquid medications, skin patches, eye drops, creams, salves, inhalers (puffers), and injections, as well as cold or allergy medications, vitamins, herbal remedies, and other supplements. The letter you received about this appointment included a plastic medications bag for all your current medications and asked you to bring them to the clinic. Have you brought this bag with you? Are these all the medications that you have taken in the past two weeks?

If “yes”, ask to see the medications and fill in the boxes in Section B.

If “no”, make arrangements to obtain medications at another time but continue with Section B and fill in the boxes as best possible by interviewing the participant for the information.

If “refused”, record reason for refusal in Comments Section

If “took no medicines”, form is complete.

1. 	Medication containers may be unavailable to the interviewer for a variety of reasons. Regardless of the reason, however, the interviewer should make an attempt to obtain the information necessary to complete the medication form.

2.	If the participant forgets to bring medication(s) to the clinic, the interviewer is responsible for obtaining the necessary information at a second visit or by telephone interview.

3.	If the participant remembers additional medication(s) taken during the previous two weeks, the interviewer should record as much information about the medication as possible at the time of the visit and then follow up with a telephone interview to check for accuracy and completeness.

4.	If the medication containers are unavailable because the participant refuses to bring them to the clinic, the interviewer should document the reason for refusal in the Comment Section. The interviewer should then attempt to obtain the participant’s cooperation in obtaining the data, either by a second visit or by telephone.

5.	If the participant brings a list of medications, instead of the medication containers, record all pertinent information from the list and note this in the Comments Section. If the interviewer has any doubt about the accuracy of the list, a follow-up telephone call should be scheduled to confirm what has been recorded.

Section B.	Prescription Medications

1.	The interviewer transcribes the name and dosage information from each medication container onto the Medications Form using the following guidelines:

2.	Medication name. Print complete medication name using block capital letters. Record all characters and numbers referring to strength as well as the units. The name of each medication should be recorded exactly as it is written on the container. Medication names that are misspelled or otherwise recorded incorrectly will cause data entry and analysis problems because they will not match the drug database. Do not record flavors of products or whether the preparations are sugar-free or sodium-free. If the medication name is longer than the 20 spaces available on the form, transcribe as much as possible and then record the complete medication name in the Comments Section. If it is not possible to transcribe the medication name, insert an asterisk (*) and explain in the Comments Section.

3.	Combination Medications contain two or more drugs. Some combination medicines, such as Dyazide, come in only one fixed combination (hydrochlorothiazide 25mg and triamterene 50mg). These combination medicines do not usually list strength. Record the name in the “Medication Name” space and leave the “Strength” column blank.

Other combination medications are available in more than one fixed dose combination. For example, Inderide, which is a combination of propranolol and hydrochlorothiazide, is available as propranolol 40mg and hydrochlorothiazide 25mg, or propanolol 80mg and hydrochlorothiazide 25mg. These combination medications usually list the strength as in “Inderide 40/25" or “Inderide 80/25." For these medications, record the name in the “Medication Name” space and the strength combination (e.g., 40/25) in the “Strength” space.

4.	Strength.

	Record the strength of each medication in milligrams (mg) whenever possible, beginning with the first space on the left in the “Strength” column.

	When strength is in milligrams, do not record the abbreviation “mg;” record only the amount of drug (e.g., if the strength is “250 mg,” record only “250”).

	When strength is not recorded as milligrams, record all numbers, digits, and characters used to denote strength, including:

- milliliter (ml)

- per milliliter (/ml)

- milliequivalent (mEq)

- hour (hr)

- per hour (/hr)

- percent (%)

	When strength is separated by a “/” (e.g. 40/25, as in combination medications), record them in this section.

	When strength is given in grains (gr), convert to milligrams using the following formula: (number of grains) x 65 = number of milligrams. (1 gr = 65 mg.)

	When strength is given in micrograms (mcg or µg), convert to milligrams using the following formula: (number of micrograms)  1000 = number of milligrams. (1000 mcg = 1 mg.)

	When strength is given in milligrams per milliliter (mg/ml), as is often the case with liquid medicine, record as in the following example: Ampicillin 125 mg / 5 ml is recorded as “125/5 ml.” (Note omission of “mg.”)

	When strength is given as a percentage (%), record as such.

	When strength is given in units (U) or units/milliliter (U/ml), as is often the case with Insulin, record as in the following examples:  “100/ml” or “100U/ml.”

	When it is not possible to record the strength, such as when it is not recorded on the medication label, record an asterisk (*) and explain in the Comments Section.

	Note: Do not record in the “Strength” column the number or quantity of medication items (e.g., number of tablets or tablespoons). See “Number Prescribed,” below.

5.	Number Prescribed. This column is designed to capture information on the number of pills (or milliliters, drops, units, etc) prescribed as opposed to the number actually taken.  Information on the number prescribed should be taken from the medication labels.

	Record the total number of medication items (e.g., “tablets”) prescribed per the given time period (e.g., day, week, or month). Circle the appropriate letter in the “Number Prescribed” column to show whether the prescribed number is per day (D), per week (W), or per month (M).

	If the instructions include a range in the number of medication items and/or times/day (or week or month) they are to be taken, record the lowest number of each. For example, if the label says, “take 1–2 tablets 3–4 times per day,” record as “3 tablets/day” (i.e., 1 tablet 3 times/day = 3 tablets/day); or, if the label says, “take 1–2 tablets every 4 hours while awake,” record as “5 tablets/day” (i.e., 1 tablet every 4 hours from 7 a.m. to 11 p.m.).

	When it is not possible to record the number of medication items prescribed per day, record an asterisk (*) and explain in the Comments Section.

	When instructions read “take as directed,” record “1” as the number prescribed per day.

	When dosing instructions are complex (e.g., “take 1 pill every other day, alternating with 2 pills every other day”), record the average number per day (or week or month).

6.	Number Prescribed: Specific Medications.

	Pill/Tablets/Capsules: Record the total number prescribed per day (or week or month).

	Solutions: Record the total number of milliliters prescribed per day (or week or month). Use the following conversions:

- 1 teaspoon = 5 ml

- 1 tablespoon = 15 ml

- 1 ounce  = 30 ml

	Eye Drops: Record the total number of drops prescribed per day (or week or month). For example, “two drops in right eye, three times a day” = 6 drops, or “one drop in each eye, twice a day” = 4 drops.

	Inhalers (puffers): Record the total number of sprays or puffs prescribed per day (or week or month).

	Insulin: Record the total number of units injected per day  (or week or month).

	Creams/Lotions/Ointments: Record the total number of applications prescribed per day (or week or month).

	Patches: Record the total number to be applied to the skin per day (or week or month).

	Nitroglycerin Ointment: Record the total number of inches to be applied to the skin per day (or week or month).

7.	PRN (“as needed”) Medication is generally used for allergy, pain, or sleep; sublingual nitroglycerin is also used PRN.

	Use the “PRN Medicine?” column to indicate whether the medication is prescribed to be taken on an “as needed” basis.

	Circle “Y” only when the prescription instructions state “as needed,” “when needed,” “if needed,” etc.

	Circle “N” when the prescription instructions do not use the words “as needed,” “when needed,” “if needed,” etc.

	The words “as directed” do not mean the same as “as needed.”

8.	Number Taken. This column is designed to capture information on the number of pills (or milliliters, drops, units, etc) actually taken as opposed to the number prescribed.  Information on the number actually taken should come directly from the participant.  People do not always take their medications as prescribed.  It is important to record information about both the number prescribed and the number actually taken as accurately as possible.

	Ask the participant, “On average during the last two weeks, how many of these pills (or other medication items) did you take per day (or week or month).”

	Record the average number of pills (or other medication items) taken per day (or week or month) during the last two weeks.

	Code “0” if none of the medication items was taken during the previous two weeks. This includes instances in which a prescription was filled but none of the medication was taken during the past 2 weeks.

	When the number taken cannot be determined, record two asterisks (**) and explain in the Comments Section.

	Circle the appropriate letter (D, W, M) to show whether the prescribed medication was taken per day, per week, or per month.

C.	Over-the-Counter Medications

Complete this section following instructions for Section B, above, but disregarding the instructions pertaining to “Number Prescribed” and “PRN Medication.”

D.	Chinese and Other Traditional Medicines

Whenever possible, in the comment section, record traditional medicine use in the same fashion as with other medicine i.e. name, dosage, frequency. If this is not possible, record the purpose of the medicine.

3.4.6  	Physical Activity

I.	Purpose

The MESA Typical Week Physical Activity Survey (TWPAS) is designed to identify the time and frequency spent in various physical activities during a typical week in the past month.  The rationale for the selected time frame of a typical week in the past month is the intention to capture typical activity patterns in one’s daily life.

The survey has 28 question items in categories of household chores, lawn/yard/garden/farm, care of children/adults, transportation, walking (not at work), dancing and sport activities, conditioning activities, leisure activities, occupational activities and volunteer activities.

II.	Methods

1.	General Instructions

1.1	This is an interviewer-administered questionnaire. Questions should be read to the participant verbatim as they appear on the form to ensure standardization.  In addition, any introductory and transitional wording should be read verbatim.

  - For the majority of the questions, response choices are “Yes” or “No”.  Record the response given by the participant. If “Yes”, ask for the number of days per week, amount of time in hours and/or minutes per day the participant spent doing the particular activity.

2.	Specific Instructions

Items to be completed by the interviewer:

	For Date, enter date form was completed. For example, September 7, 2002 would be entered as 09/07/2002.

	ID# & Acrostic will be pre-printed.

- Begin the questionnaire by reading the following instructions to the participant:

Think about the types of activities you did in a typical week in the past month. Please indicate whether you did or did not perform each of the following activities in a typical week.  For each item that you respond ‘yes,’ you will be asked for the number of days in a typical week you did these activities and the average amount of time per day in hours and minutes.

Define intensity levels for the participant:

Most of the survey questions ask about light, moderate, and heavy intensity activities.

Light intensity refers to activities that require little effort and are easy to do.

Moderate intensity refers to an effort that is harder than light intensity but is not an all-out effort. Vigorous intensity is a very hard activity and requires all-out effort.

Show the participant using the following example: Begin this with, “Let me show you an example of how we will fill out the survey.”

To orient him/her to the past month, you will identify that period for the participant. In the text box above, if we assume, for instance, that the current date is September 15, the past month would start on August 15.

Using the example below and review each step with the participant. Give him/her time to consider each step and to ask questions. Explain that, if the participant continued usual physical activities while on vacation (or during some other atypical period of time), he/she should report them as usual for a typical week. However, if usual activities were stopped during a vacation, or if the participant took up other activities only during that period¸ then he/she should not record them as typical activities.

Questions people might ask about completing the TWPAS, and some sample responses:

	What if every week in the past month was different?

Think about the week that was most typical of your activity patterns for that activity in other times of the year and fill in the circle for the number of days and hours per day and or minutes per day.

	What if the length of time is different each day?

Think about the average in all the days that reflects your typical time for the activity in a typical week.

	I was on vacation in the past month when I went on a 2-week bicycle trip.  Should I include this in the estimates?

In this case, think about a typical week in the past couple of months that reflects your usual activity patterns.  We are trying to identify the activity patterns you do on a usual basis, so if your vacation was not typical, do not include it.

Once you have reviewed the sample question with the participant and explained the difference between typical and atypical activities, ask if he/she has any questions.

It is possible that a person will spend more time doing activities on one day or another (e.g., weekends).  If this is the case, ask him/her to estimate the usual time during each event in a typical week, averaging in the longer and shorter days.

For example, if the participant engages in an activity for 30 minutes/day 5 days/week and 2 hours/day 1 day/week, ask him/her to add about 15 minutes extra to each day, to take into account a single day that has a prolonged bout of activity compared to the usual.

3.	Item-by-Item Clarification

Household chores

“In a typical week in the past month, did you do… (light effort household chores).” The examples should be read to the participant for clarification.

If “no”, bubble in the no and proceed to the next question.

If “yes”, then ask, “About how many days per week?” bubble in the number of days.

Then ask, “About how many hours or minutes per day? Bubble in the number of hours and/or minutes e.g. the participant response 1hr and 15 mins,  bubble in 1 under hours and 15 under minutes. These series of questions should be applied to similar types of questions in the questionnaire.

1.	Light effort Household Chores:  These activities are light intensity, routine, usually daily activities that people do during the care and maintenance of a household. Examples include cooking and cleaning after cooking, straightening up the house, grocery and household shopping and putting things away, changing the bed, doing the laundry, ironing. Housecleaning in a structured, organized way should not be included here, as that would involve more moderate intensity chores.

2.	Moderate or Heavy Effort Household Chores:  These activities are more structured and might not occur on a daily basis. Examples include heavy cleaning (washing windows, moving furniture to clean), vacuuming, scrubbing the floors or walls, mopping—either standing up or on hands and knees—repairing home appliances or lawn and garden tools, washing the car.

Lawn/Yard/Garden/Farm

“In a typical week in the past month, did you do… (moderate effort lawn/yard etc.)”

3.	Moderate Effort Lawn/Yard/Garden/Farm etc. activities:  These activities refer to outside chores involved in caring for a house, farm, or ranch.  They may involve yard work, cleaning out the garage, raking the leaves, sweeping the porch or sidewalk, or other moderate effort chores.  Encourage the participant to think of activities done in a typical week in the past month. This category may include seasonal activities; if so, the activities reported should be typical of the past month.

4.	Heavy Effort Lawn/Yard/Garden/Farm etc. activities:  These activities require heavy effort and may be seasonal. Examples include digging dirt, shoveling snow or using a snow blower, chipping ice, tilling a garden, chopping and hauling wood, and removing trees.

Care of Children/Adults

“In a typical week in the past month, did you do… (light effort child/adult care)”

5.	Light Effort Children/Adult Care: These activities require physical movement by the respondent and include bathing, feeding, changing diapers, playing with a child, or other similar activities. Do not count time sitting with a child (as in babysitting) without active engagement in physical activities. Include only the time spent involved in physical activities.

6.	Moderate Effort Children/Adult Care: These are intentional activities that require moderate effort to complete and may include activities of lifting and carrying dependent others, pushing a wheelchair or stroller. Include only the time spent moving.

Transportation

“In a typical week in the past month, did you … (use transportations)”

7.	Drive or ride in a car, a bus, subway etc as a mode of transportation. These are all light activities and do not include walking. Have the respondent include all time spent in transportation where they are sitting in any vehicle above.

Walking (not at work)

“In a typical week in the past month, did you do… (walking to get to places, etc.)”

8.	Walking to get places.  In general, walking is underreported in the time estimates. This would include walking for transportation, walking to and from work, walking to the store or from the car into the store and back, walking to get the mail, etc.

9.	Walking for exercise, pleasure, social reasons, walking during work breaks, or walking the dog. These are classified as intentional walking and are different in intention then those in item #8. The walking may be for exercise or part of a daily routine that is done with family members, animals, or for personal reasons. Walking for transportation should be included in item 8.

Dancing/Sport Activities

“In a typical week in the past month, did you do… (dancing in church/team sport, etc.)”

10.	Dancing in church, ceremonies, or for pleasure. Remind the participant to think of a typical week in the past month to estimate usual dancing behaviors. Some may dance only occasionally—a few times a year. This would not be included as a “yes” response to this category unless it was typical of the past month. Ceremonial or religious dancing would need to be done regularly enough to represent a typical week in the past month.

11.	Team sport. The purpose of including team sports activities is to group exercise activities that are done with others. These are probably seasonal activities that are done in leagues or other organized settings.  Remind the participant to think of a typical week in the past month and to stay within that framework when responding.

12.	Dual sports. These activities involve mostly racket sports or other one-on-one sports activities. They could include fencing, ping-pong, or other activities done with another person.

13.	Individual activities that maybe classified as sports. These may be sports activities, such as golf and bowling, or more individual relaxation/meditation activities, such as yoga or Tai Chi. Remind respondents to think of days and time spent during a typical week in the past month only.

Conditioning Activities.

“In a typical week in the past month, did you do… (moderate effort conditioning, etc. )”

14.	Moderate Effort Conditioning Activities. Conditioning activities are those that can be done alone or with others. They are different from sports, because the intention is to gain an element of fitness rather than have a contest or win a game. Moderate effort activities are not for competition, nor are they all-out effort. Intensity of exercise should be moderate enough that respondents should be able to talk with others while they are performing the activities. Examples are low impact aerobics, recreational (slow) bicycling, rowing on a rowing machine or in a lake, swimming in a pool or lake, or using weight lifting or conditioning machines at a health club.

15.	Heavy Effort Conditioning Activities. These are very intense activities done for maximum fitness levels and include high impact aerobics (e.g., Tai-bo, kick boxing), competitive or maximum effort running, bicycling, swimming, and work on health club machines. Exercise at this intensity would be very hard and the respondent would have difficulty carrying on a conversation during the performance.

Leisure Activities.

“In a typical week in the past month, did you do…(these leisure activities, watching TV)”

16. 	Watching TV.  This is a sedentary, leisure-time pursuit. Do not include the time watching TV while doing other things. The question is to be used as a marker of sitting or reclining and watching TV as a single pursuit.

17.	Read, knit, sew, visit, do nothing, non-work recreational computer.  These are markers of sedentary activities that exclude watching TV.  If people report knitting while watching TV, make sure they select one category or the other and avoid duplicate category reporting.

Occupational Activities.

18.	Do you work to earn money?

If “no,” skip to Volunteer Activities, question #24.

If “yes,” continue to question #19.

Respondents should report activities only if they work to earn money.  Volunteer activities will be asked about next. Unpaid household activities and caring for others were included in questions #1–6.

19.	How many days/week and hours/day worked. If the respondent works at more than one job, record the total days worked per week in all jobs. To estimate the average hours worked per week, ask the respondent to consider the hours they work in the second job (or third, if applicable) and then add them to the average time they work per day. See the example about averaging time per day when activities are done longer on some days of the week than on others.

Respondent should fill in the circles for the time spent in each activity at work. The hours per day for all activities should equal the total hours per day worked. Recommend that the respondent keep a cumulative total of the hours recorded doing occupational activities, in order to avoid estimating more time in occupational tasks than time they work on an average day. They can also subtract the hours from the total time worked per day as they identify how the time is spent in different types of activities.

“At work, did you do….”

20.	Light Effort/Sitting Activities. For most respondents, this will be the most hours in the work day.

21.	Light Effort/Standing Activities. These are likely intermittent activities that would be done in a clerical setting (e.g., office work related to filing, using a copy machine) or sustained activities done in a labor setting (e.g., check-out clerk in a store, assembly line worker assembling parts, medical field examining patients). Teaching in a classroom falls into this category.

22.	Moderate Effort/Standing or Walking Activities: For some occupations (office work, clerical, professional), these may be more intermittent, as in walking down the hall, walking between office buildings, and delivering items. For labor settings, this could relate to jobs such as delivery person (overnight express delivery, food delivery, mail delivery) or jobs that require mostly walking and standing (nurse, custodian, physical education teacher, coach, firefighter, police officer, physical therapist).

23.	Heavy Effort/Manual Labor: These occupations require manual effort that involves substantial movement and labor. Types of activities may include digging ditches, ranch or farm labor, delivering furniture, loading and unloading trucks, seasonal farm work.

Volunteer Activities

24.	Did you work as a volunteer and/or work at church doing activities you have not yet mentioned on this survey?

If “no,” skip to question 28.

If “yes,” continue to question 25.

This question is designed to identify time spent in unpaid, work-like activities outside the home.

Did your volunteer work include….

25.	Light Effort Activities. The respondent should identify light activities that required little effort and were done predominantly while sitting or standing (e.g., cooking at church banquet).

26.	Moderate Effort Activities.  The respondent should identify activities that required moderate effort and were done predominantly while standing and walking (e.g., participating in a neighborhood clean-up project).

27.	Heavy Effort Activities. The respondent should identify activities that required heavy effort, such as pushing, lifting, carrying, and climbing (e.g., building a house for Habitat for Humanity.

Walking Pace

28.	When you walk outside of your home, what is your usual pace?

Ask respondent to estimate the usual pace he/she walks most of the time. Consider all walking activities (e.g., at work, on the way to work, for exercise, in walking with children or others, or when running errands). Fill in appropriate circle.

Offer the following guidelines:

	Slow or Casual strolling pace = 2 mph = 30 minutes per mile

	Average or normal pace = 2-3 mph = 20-30 minutes per mile

	Fairly brisk pace = 4-5 mph = 12-15 minutes per mile (very fast or almost a slow jog)

	Brisk or striding pace = More than 5 mph = 10 minutes per mile (race-walking)

After completion of the interview, the interviewer should check to make sure all questions were answered then fill in the Interviewer ID at the end of the survey.

A reviewer should then review the form for completeness, verify any information that seems unclear and attempt to complete any thing missing by asking the interviewer and participant about specific skipped questions. Then complete the questionnaire by filling in the reviewer ID box.

3.4.7  	Health and Life

I.	Purpose

This questionnaire includes several instruments designed to measure psychosocial characteristics that may be important in understanding the causes of cardiovascular disease. These psychosocial factors may themselves lead to increased risk of cardiovascular disease or may interact with other traditional risk factors, such as diet or sedentary lifestyle. The areas assessed as part of this questionnaire include temperament; attitude toward health, job, relationships; feelings of happiness/unhappiness; availability of help/affection. Each area is measured by a set of questions or a scale.

II. 	Definitions

The terms used in the questionnaire should require no explanation, because they are used in the way they tend to be used by most people in everyday life.

III. 	Methods

This is a self-administered questionnaire. Provide the participant with the form and a pencil and give brief instructions for completion.

1.	General Instructions

1.1	It is important that the participant have some private time in a quiet area to complete the form. The participant should be asked to answer each question by bubbling in the circle with the appropriate response. Review the top section of the questionnaire carefully with the participant before starting. Emphasize that there are no right or wrong answers and that we are interested in their feelings and opinions about things. Also emphasize that they should not spend too much time on any one question.  Show them that additional instructions are provided at the beginning of each section. Before starting, note that only one bubble should be filled for each question or statement.

1.2	Ask the participant to try to answer all questions. Remind him/her to request assistance from a staff member if anything is unclear. Most participants should be able to complete the questionnaire on their own. However, if the participant expresses or appears to have difficulty reading or comprehending the questions, offer to help and make arrangements for an interviewer-administered version in the appropriate language.

1.3	Important points for clinic staff and participants to consider:

	If the topic should arise, remind participants that all information is strictly confidential and will be used only for research purposes. Explain that things about people’s lives, including the stressful situations they go though, may be important to their health. Knowing about these things may help us understand the causes of heart disease better. Also emphasize that it is important to get complete data so that the study results will be valid. However, if a participant is upset by the questions or does not want to answer, he or she should feel free to skip the question or section. Refusal to answer the questions will not jeopardize his/her participation in the study.

	The measurement of these dimensions is complex. Generally they are measured using scales or collections of questions that attempt to get at the same underlying concept in different ways. For this reason some of the questions may seem repetitive. If questions on this should arise, acknowledge that some questions may seem similar, but ask participants to respond to each one separately as best they can.

	The terms used should be understood by most people. If the participant asks about the meaning of any item or tries to qualify a statement, ask the participant to re-read the statement (or question) and answer as they best understand the question. Do not attempt to explain the question or provide synonyms (unless specified in the specific instructions below), because this may create problems for some of the scales.

2.	Specific Instructions

Question 1:  Are you currently married or living with a partner?

This question requires a “yes” or “no” answer.  The participant should answer “yes” if s/he is married or living with a partner.  If the participant is married or has a partner but is currently not living with that person, then “no” should be marked.

Questions 2A-2T correspond to a scale designed to measure a person's temperament. The participant is asked to indicate how well each statement describes the way he or she usually is.  The instruction on the form reads:

For each of the following statements, please choose the one response that best describes you.

Possible answers are: “Almost never,” “Sometimes,” “Often,” “Almost Always.”

- A.  I am quick tempered.

- B.  I have a fiery temper.

- C.  I am a hotheaded person.

- D.  I get angry when I’m slowed down by others’ mistake.

- E.  I feel annoyed when I am not given recognition for doing good work.

- F.  I fly off the handle.

- G.  When I get mad, I say nasty things.

- H.  It makes me furious when I am criticized in front of others.

- I.  When I get frustrated, I feel like hitting someone.

- J.  I feel infuriated when I do a good job and get a poor evaluation.

- K.  I am a steady person.

- L.  I feel satisfied with myself.

- M.  I feel nervous and restless.

- N.  I wish I could be as happy as other seem to be.

- O.  I feel like a failure.

- P.  I get in a state of turmoil or tension as I think over my recent concerns and interests.

- Q.  I feel secure.

- R.  I lack self-confidence.

- S.  I feel inadequate.

- T.  I worry too much over something that does not matter.

Questions 3-7 correspond to a scale designed to measure general quality of life.

This section begins with this general statement:

Many people experience ongoing problems in their everyday lives.  Please tell us whether any of the following has been a problem for you.

For each question, in part A the participant is asked whether s/he has experienced the problem or situation identified (“yes” or “no”).  If the participant answers “yes,” then the participant is instructed to indicate in part B whether the difficulty has lasted for six months or longer (“yes” or “no”) and in part C whether s/he would say the problem has been “Not very stressful,” “moderately stressful,” or “Very stressful.”

3.  Serious ongoing health problem (yourself).

4.  Serious ongoing health problem (someone close to you).

5.  Ongoing difficulties with your job or ability to work.

6.  Ongoing financial stress.

7.  Ongoing difficulties in a relationship with someone close to you.

Questions 8A-8T correspond to a scale designed to measure a person's feelings and behavior during the past week. The participant is asked to indicate how well each statement describes his/her feelings or behavior during the past week.  The instruction on the form reads:

Below is a list of the ways you might have felt or behaved.  Please indicate how often you felt this way DURING THE PAST WEEK.

Possible answers are:	“Rarely or none of the time (Less than 1 Day),”

“Some or a little of the time (1-2 Days),”

“A moderate amount of the time (3-4 Days),”

“Most of the time (5-7 days).”

- A.  I was bothered by things that don't usually bother me.

- B.  I did not feel like eating; my appetite was poor.

- C.  I felt that I could not shake off the blues, even with help from my family and friends.

- D.  I felt that I was just as good as other people.

- E.  I had trouble keeping my mind on what I was doing.

- F.  I felt depressed.

- G.  I felt that everything I did was an effort.

- H.  I felt hopeful about the future.

- I.  I thought my life had been a failure.

- J.  I felt fearful.

- K. My sleep was restless.

- L.  I was happy.

- M.  I talked less than usual.

- N.  I felt lonely.

- O. People were unfriendly.

- P.  I enjoyed life.

- Q.  I had crying spells.

- R.  I felt sad.

- S.  I felt that people dislike me.

- T.  I could not “get going.”

Questions 9-14 correspond to a scale designed to measure a person's feelings of isolation or support.  The participant is asked to indicate how well each statement describes the way he or she usually is.  The instruction on the form reads:

Please read the following questions and mark the answer that best describes your life now.

Possible answers are: “None of the time,” “A little of the time,” “Some of the time,” “Most of the time,” and “All of the time.”

9.  Is there someone available to you whom you can count on to listen to you when you need to talk?

10.  Is there someone available to give you good advice about a problem?

11.  Is there someone available to you who shows you love and affection?

12.  Is there someone available to help you with daily chores?

13.  Can you count on anyone to provide you with emotional support (talking over problems or helping you make a difficult decision)?

14.  Do you have as much contact as you would like with someone you feel close to, someone in whom you can trust and confide?

At this point the participant questionnaire section is complete.

MESA staff will complete the form by filling out the box “For MESA Field Center Use Only:”

- If the form was self-administered, check for completeness.

- Mark if form was self-administered or interviewer-administered.

- Record Interviewer or Reviewer ID (your ID number)

- Record Data Entry ID.

3.4.8  	Family History

I.	Purpose.

The Family History is designed to gather information regarding the health history of the participant’s biological (natural) parents and biological (natural) siblings (including all blood related siblings e.g. half-siblings).  Detailed family histories provide important information regarding familial clustering of disease.  Participant reports of family histories have been employed in several family studies (i.e. the NHLBI Family Heart Study (FHS), the Family Blood Pressure Program) with good success. The MESA study is particularly interested in family history of medical conditions that are related to atherosclerotic (plaque) formation and diseases as well as cancer and asthma.  Specifically, we are interested in family history of heart attack, stroke, high blood pressure, diabetes, and other heart diseases, i.e. valvular heart diseases, septal heart defects etc.

II.	Definitions

In this questionnaire, an attempt has been made to use common terms that require no explanation, because they are used in the way they tend to be used by most people in every day life.  However, some terms are more scientific and may require clarification.  In such cases, please refer to the Definition and Synonyms for Medical Terminology at the end of the instruction.

However, it is critical that the participant understand that we are interested in gathering information only about his/her biological or natural parents and biological siblings.  It should be explained that full biological siblings are brothers and sisters who have the same biological or natural parents as the participant.  If indicated, it can additionally be explained that biological family members differ from adopted, step, or foster family members in that they have a genetic or “blood” relationship to the participant.

III.	Methods

1.	General Instructions

This is an interviewer-administered form. Questions should be read to the participant verbatim as they appear on the form to ensure standardization.  In addition, any introductory and transitional wording should be read verbatim.

If the topic should arise, remind the participant that all information is strictly confidential and will be used only for research purposes. Explain that information about the Family History may be important in understanding their health. The information may, in turn, help us to better understand the causes of heart disease.

2.	Specific Instructions for Completing the Family History Questionnaire

- The interviewer should take time to become familiar with the Family History form.  Note that the form is arranged with the disease of interest in the columns along the top (x-axis) and the “blood” relation of interest is in rows down the page (y-axis).  Proceed with the first “relation” and ask about the diseases of interest pertaining to that particular “relation” along the row and across the page.  Response choices are: Yes (Y), No (No), or Don’t know (D).

For example: First begin with “Spouse”, skip blacked out boxes and proceed to the next column.  In this case, under the column “Age at last birthday or age at death,” fill in the digits for age of last birthday or at death.  Then, under the next column “Heart Attack or M.I.” fill in “Y, N, or D”.  If yes, ask for the age of first heart attack; fill in the digits for the age then continue to the next column.  For each column of disease or condition, if the response is “Yes”, ask for the age of when this first occurred.  For the column asking about cancer, in addition to age, ask the participant for the type of cancer.  If the participant responded with more than one cancer, record the cancer that was diagnosed first.

- Begin the questionnaire with the first column and the first row – regarding “Spouse” by asking:

Do you currently have a spouse?

If “no”, skip to the next “relation-person”, mother.

If “yes”, the interviewer should say, “the next series of questions are about your spouse.”

Then proceed with the next column:

Definition: Spouse is defined as a married person by state or common law.  For the purpose of this questionnaire, same sex domestic partnership will not be considered as spouse.

What was your spouse’s age at the last birthday? If the spouse is deceased, then ask, “What was your spouse’s age at death, (when he passed away)?

Round off the age to the nearest whole number and enter the digits into the boxes then proceed to the next column of disease of interest.  If there are more than one deceased spouse, only collect information on the most recent spouse.

- Has your spouse ever had a heart attack or Myocardial Infarction?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “at what age did your spouse have the first heart attack?”

Enter the age at first heart attack into the boxes then proceed to the next column.

- Definition: To aid the participant, a heart attack can be defined as damage to, or death of, part of the heart muscle due to insufficient blood supply.

- The intent is to record whether the first occurrence of one of these events indicates that coronary heart disease was before or after a certain age.  If the participant cannot answer this question confidently, the response should be coded as “Don’t Know."

- Has your spouse ever had heart surgery or a cardiac (heart) procedure, such as coronary bypass surgery, balloon angioplasty, or stent placement because arteries of the heart were blocked?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “at what age did your spouse have this procedure?”

Enter the age when the first procedure was done into the boxes then go to the next column.

Definition: Coronary bypass can be defined as surgery on the heart to improve blood supply/flow to the heart muscle.  This surgery is most often performed when narrowed coronary arteries reduce the flow of oxygen-containing blood to the heart.  Balloon angioplasty can be defined as a procedure to dilate (widen) narrowed coronary arteries.  A catheter with a deflated balloon on its tip is passed into the narrowed artery segment, the balloon is inflated, and the narrowed segment is widened.  Acceptable symptoms for heart attack, coronary bypass surgery, or balloon dilation are listed in the Definition and Synonyms for Medical Terminology at the end of these instructions.

The intent is to record whether the first occurrence of one of these procedures indicating coronary heart disease was before or after a certain age.  If the participant cannot answer this question confidently, the response should be coded as “Don’t Know."

- Has your spouse ever had a stroke?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “at what age did your spouse have the first stroke”.

Enter the age at first stroke into the boxes then proceed to the next column.

Definition: To aid the participant, a stroke can be defined as the sudden loss of muscle function, vision, sensation, or speech resulting from brain cell damage caused by insufficient supply of blood to a part of the brain.  A stroke may result from blockage of the arteries by blood clots, atherosclerotic plaques, or rupture of an artery.  Cerebral hemorrhage can be defined as bleeding into or around the brain caused by a rupture of an artery or aneurysm in the brain.  Cerebral hemorrhage usually causes sudden onset of very severe headache and is associated with other symptoms described above for stroke or with loss of consciousness.  DO NOT include TIAs or ‘mini-strokes’.  Acceptable synonyms for stroke and cerebral hemorrhage are listed in the Definition and Synonyms for Medical Terminology at the end of this instruction.

The intent is to determine whether the first time one of these events occurred prior to 60 years of age.  If the participant cannot answer this question confidently, the response should be coded as “Don’t Know."

- Does/Did your spouse have high blood pressure or hypertension diagnosed by a physician?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “at what age was your spouse first diagnosed with high blood pressure?”

Enter the age at first diagnosis of high blood pressure into the boxes then proceed to the next column.

A  “Yes” response means that, to the best of the participant’s knowledge, a physician specifically told his/her spouse of the diagnosis of high blood pressure or hypertension.  Other conditions, such as being anxious, tense, or “hyper” should not be confused with high blood pressure.

Definition: To aid the participant, the interviewer can explain that blood pressure is measured using a device similar to the one used in this study.  Based on those measurements, a physician may have diagnosed the person as having high blood pressure or hypertension.  A spouse who may have made a self-diagnosis based on readings obtained using an automatic device in a mall or other location should not be considered as having high blood pressure, unless a doctor also verified the condition.  In most cases, unless a doctor has prescribed anti-hypertensive medication or arranged a plan for follow-up blood pressure readings, a diagnosis of high blood pressure is in doubt.

Does/did your spouse have Diabetes or high blood sugar diagnosed by a physician?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “at what age was your spouse first diagnosed with Diabetes?”

Enter the age at first diagnosed of diabetes into the boxes then proceed to the next column.

A “Yes” response to the question means that a physician has specifically told the participant’s spouse that he/she has diabetes or high blood sugar.  Other blood sugar conditions, such as “hypoglycemia”, should not be considered a “Yes” response to this question.  If a physician had told the participant’s spouse that he/she has “borderline” or “early” diabetes (or high blood sugar), code the response as “yes”.

Does/did your spouse have Cancer diagnosed by a physician?

- If “no” or “don’t know”, fill in the appropriated circle then skip to the next column.

If “yes”, ask, “was this a cancer other than skin cancer such as basal cell or squamous cell skin cancer?”

A “Yes” response to the question means that a physician has specifically told the participant’s spouse that he/she had cancer.  An attempt should be made to encourage the participant to specify the type of cancer.  If this is not possible, leave the fill in box blank.

If this cancer was a type other then the two mentioned then continue below; otherwise skip to the next column of disease/condition:

“Do you know the specific type of the cancer your spouse had?”

“At what age was your spouse diagnosed with this cancer (or the very first cancer if more the participant gives more than one cancer)?”

Enter the age when first diagnosed and the specific type of cancer into the appropriate boxes then proceed to the next column of disease of interest.

- Does/did your spouse have any other heart diseases/problems diagnosed by a physician i.e. heart valve problems, heart defects, heart enlargements, heart conduction disorders or irregular rhythm disorders?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “At what age was your spouse first diagnosed with this heart disease/problem?”  Enter the age at first diagnosed into the boxes then proceed to the next column.

A “Yes” response to the question means that a physician has specifically told the participant’s spouse that he/she had some other heart disease/problems other than those specifically asked previously.  An attempt should be made to encourage the participant to specify type of heart disease/problem.  This will ensure that indeed some other heart disease exists.

- Does/did your spouse have Asthma diagnosed by a physician?

- If “no” or “don’t know”, fill in the appropriated circle then skip to the next column.

If “yes”, ask, “At what age was your spouse first diagnosed with Asthma?”

Enter the age at first diagnosed with asthma into the boxes.

A “Yes” response to the question means that a physician has specifically told the participant’s spouse that he/she had Asthma.  It can be explained to the participant that a person with asthma usually has an inhaler that is used to improve breathing.

The interviewer has now completed all questions in the first row regarding “Spouse”.  The interviewer should now proceed to the second row.

Questions regarding the Biological/Birth Mother:

The following questions are regarding your Biological or Birth Mother, do you have any comments before we begin?

If the participant was adopted or for some reason does not have information about his/her biological/birth mother, the interviewer should confirm whether the participant responded “no,” or “don’t know” and fill in the appropriate circle.

If “no”, skip to the next row with question about the biological father.

If “yes” or “don’t know”, the interviewer will continue with the questions across the column i.e.

What was your mother’s age at the last birthday (if deceased, age at death)?

Round off to the nearest whole number and enter the digits into the boxes then proceed to the next column.

- Has/Did your mother ever had a heart attack or Myocardial Infarction?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “at what age did your mother have the first heart attack?”

Enter the age at first heart attack into the boxes then proceed to the next column.

The interviewer will continue in this fashion as done for the first set of questions until the last question regarding Asthma.  Then, the interviewer will begin with the next “Relation” column.

Questions regarding the Biological/Blood Father.

The interviewer should begin with:

The following questions are regarding your Biological Father; do you have any comments before we begin?

If the participant was adopted or for some reason does not have information about his/her biological/birth father, the interviewer should confirm whether the participant responded “no,” or “don’t know” and fill in the appropriate circle.

If “no”, skip to the next row with question about the biological siblings.

If “yes” or “don’t know”, the interviewer will continue with the questions across the column i.e.

What was your Father’s age at the last birthday (if deceased, age at death).

Round off to the nearest whole number and enter the digits into the boxes then proceed to the next column of disease of interest.

- Has/Did your Father ever had a heart attack or Myocardial Infarction?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “at what age did your Father have the first heart attack?”

Enter the age at first heart attack into the boxes then proceed to the next column.

The interviewer will continue in this fashion as done for the first set of questions until the last question regarding Asthma.  Then, the interviewer will begin with the next “Relation” column.

Questions regarding the Biological/Blood Siblings

The interviewer should begin with:

The following questions are regarding your Biological Siblings; we will begin with the oldest sibling and work our way to the youngest.  Do you have any comments before we begin?

If the participant has no siblings, skip this and proceed to the next “Relation” category – Child.

If the participant has more siblings than available on the form, use a second form to obtain complete information on all siblings.

If the participant for some reason does not have information about his/her biological/blood siblings, the interviewer should confirm whether the participant said “no,” or “don’t know” and fill in the appropriate circle.

If “no”, skip to the next row with question about the biological Children.

If “yes” or “don’t know”, the interviewer will continue with the questions across the column i.e.

What was your Sibling’s age at the last birthday (if deceased, age at death).

Round off to the nearest whole number and enter the digits into the boxes then proceed to the next column.

- Has/Did your Sibling ever had a heart attack or Myocardial Infarction?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “at what age did your Sibling have the first heart attack?”

Enter the age at first heart attack into the boxes then proceed to the next column.

The interviewer will continue in this fashion as done for the first set of questions until the last question regarding Asthma.  Then, the interviewer will begin with the next sibling or “Relation” column.

Questions regarding the Biological/Blood Children

The interviewer should begin with:

The following questions are regarding your Biological Children. We will begin with the oldest child and work our way to the youngest.  Do you have any comments before we begin?

If the participant has no children, skip this and the interview is now completed.

If the participant has more children than available on the form, use a second form to obtain complete information on all children.

If the participant for some reason does not have information about his/her biological/birth children, the interviewer should confirm whether the participant said “no,” or “don’t know” and fill in the appropriate circle.

If “no”, skip this section and the interview is now complete.  Thank the participant.

If “yes” or “don’t know”, the interviewer will continue with the questions across the column i.e.

What was your child’s age at the last birthday (if deceased, age at death).

Round off to the nearest whole number and enter the digits into the boxes then proceed to the next column.

- Has/Did your child ever had a heart attack or Myocardial Infarction?

- If “no” or “don’t know”, fill in the appropriate circle then skip to the next column.

If “yes”, ask, “at what age did your child have the first heart attack?”

Enter the age at first heart attack into the boxes then proceed to the next column.

The interviewer will continue in this fashion as done for the first set of questions until the last question regarding Asthma.  Then, the interviewer will begin with the next child.

Once all the information has been obtained for the last child, the interview is complete.  A reviewer should review the form for any incompleteness.  Then complete the questionnaire form by filling in the fields in the box labeled, “For MESA Field Center Use Only,” select whether the form was self-administered or interviewer-administered. Record reviewer (if form self-administered) or interviewer (if form interviewer-administered) ID number.

Table of Medical Terms

MEDICAL CONDITION OR PROCEDURES           DEFINITIONS & SYNONYMS

| Hypertension or high blood pressure | A chronic increase in blood pressure above its normal range that does not include being anxious, tense or "hyper". |
| --- | --- |
| Diabetes or high blood sugar | A chronic disease where the body doesn't produce or properly use insulin that occurs at times other than during pregnancy. Also called diabetes mellitus. |
| Heart attack | Death of, or damage to, part of the heart muscle due to an insufficient blood supply. Also called myocardial infarction or MI. |
| Coronary bypass surgery | Surgery to improve blood supply to he heart muscle. This surgery is most often performed when narrowed coronary arteries reduce the flow of oxygen-containing blood to the heart itself. Also called CABG, "cabbage operation"; by-pass graft or operation; vein by-pass operation; open heart surgery. |
| Balloon Angioplasty | A procedure used to dilate (widen) narrowed arteries. A catheter with a deflated balloon on its tip is passed into the narrowed artery segment, the balloon inflated, and the narrow segment widened. To keep arteries from collapsing, stents (stainless steel supports) can be inserted into the artery during angioplasty. Also called percutaneous angioplasty, balloon dilation, PTCA, stents, etc. |
| Stroke | Loss of muscle function, vision, sensation or speech resulting from brain cell damage caused by an insufficient supply of blood to part of the brain. Also called apoplexy, cerebrovascular accident, or cerebral vascular accident. |
| Transient ischemic attack, or TIA | A temporary stroke-like event that lasts for only a short time and is caused by temporarily blocked blood vessels to part of the brain. |

3.4.9	Sleep History

I. 	PURPOSE

The Sleep History Questionnaire is a means to obtain information about sleep habits, and especially about symptoms of sleep apnea (abnormal breathing during sleep).  Recent studies suggest that disordered breathing during sleep might be related to cardiovascular conditions such as hypertension, heart disease, and stroke.

II. 	Materials/Equipment

This is a self-administered form.  Please give the form and a pencil to the participant and provide brief instructions for completion.  It should take about three minutes to complete the questionnaire.

III. 	Definitions

Most of the terms on the questionnaire require no special explanation; they are used by the layperson.  The definition of “sleep apnea” (Question 8) is provided within the question.

IV. 	Methods

1.	General Instructions

1.1	The participant should complete the form privately, in a quiet room, sitting at a table, and with no sense of urgency.  Please hand the form and a pencil to the participant, and tell him/her to answer each question by darkening the circle of the appropriate response. Review the top section of the form with the participant before starting.  Please emphasize the importance of having complete and accurate information.

1.2	Ask the participant to try to respond to all questions, unless instructed to skip a question.  Show the participant an example of a skip pattern (e.g., Question 3).  Remind the participant to request assistance from a staff member if anything is unclear.  If the participant expresses or appears to have difficulty reading or comprehending the questions, offer your help and make arrangements for an interviewer-administered version in the appropriate language.

1.3	Important points for interviewers and participants to consider:

	Should the topic of confidentiality arise, please remind the participant that all collected information is strictly confidential and will only be used for research purposes.  If a participant seems upset by the questions or does not want to answer, he or she should feel free to skip the question or section. Refusal to answer the questions will not jeopardize participation in the study.

	Most people should understand the terms on the questionnaire.  If the participant asks about the meaning of any item or tries to qualify a statement, please ask the participant to re-read the statement (or question) and answer as best they understand.

2.	Specific Instructions for Completing the Sleep History Questionnaire

The questionnaire begins with the following introduction statement:

The following questions pertain to the participant’s sleeping habits.  The information will help us better understand the health consequences of sleeping disorders.  Please ask the participants to consider both what they know about themselves and what others (e.g., bed partners) have told them.

1.	Since your MESA clinic visit, have you been awakened at night by trouble breathing?

The responses include: “yes,” “no,” “not sure”.  Please select the choice that best applies then continue with question #2.

2.	Have you ever snored (now or at any time in the past)?

If “no” or “don’t know”, skip to question #5.

Otherwise, continue with the next question.

3.	How often do you snore now?

The participant should mark the response that best corresponds to the snoring frequency.

If the participant is uncertain about a choice between two adjacent categories (e.g., 3-5 nights versus 6-7 nights), instruct him/her to check the lower category.

If the response is, “Do not snore any more,” skip to question #5.

Otherwise, continue with the next question.

4.	How loud is your snoring?  Select the choice that best estimates the loudness.

5.	Are there times when you stop breathing during your sleep?

If “no” or “don’t know”, skip to question #7.

Otherwise, continue with the next question.

6.	How often do you have times when you stop breathing during your sleep?

Select the choice that best estimates the frequency.

If the participant is uncertain about a choice between two adjacent categories (e.g., 3-5 nights versus 6-7 nights), instruct him/her to check the lower category.

7.	How often do you feel excessively (overly) sleepy during the day?

Select the choice that best estimates the frequency.

If the participant is uncertain about a choice between two adjacent categories (e.g., 2-4 days/month versus 5-15 days/month), instruct him/her to check the lower category.

8.	Have you ever been told by a doctor that you had sleep apnea (a condition in which breathing stops briefly during sleep)?

Participants who are not familiar with the term “sleep apnea” should answer “No”.  If the participant heard the term before, thinks that a doctor might have told them that they had sleep apnea, but are not quite sure -- they should mark “Don’t know”.

9.	How often do you “make time” in your schedule for a regular nap or “siesta” in the afternoon?

If “never or rarely”, the questionnaire is now completed and the form should be returned to a MESA staff member.

Otherwise, continue with the next question.

10.	When you do nap in the afternoon, how long do you sleep?

The participant should respond with the best whole number of hours and/or minutes.

Participants might forget to enter “0” hours for naptime shorter than 60 minutes, might forget to enter “00” minutes, or might enter values such as “80” minutes.  If you note any type of mistake or omission, please verify the participant’s intention and correct accordingly.  (Do not assume that you understand the participant’s intention.)

11. 	What are your reasons for regular napping in the afternoon?

(Check all that apply.)  If “other”, the participant should specify the reason in the box.

At this point the participant questionnaire section is complete.

A MESA staff member will complete the form by filling our the box “For MESA Field Center Use Only:”

- If the form was self-administered, please check the form for completeness and adherence to the skip patterns.  If you note a skip-pattern error or omission of a response, please clarify with the participant.

- If the participant forgot to enter the date on the top of the form, please fill out the date.

- Mark self-administered or interviewer-administered.

- Record Interviewer or Reviewer ID (your ID number)

- Record Data Entry ID.

3.4.10	Eye History

I.	PURPOSE

The Eye History questionnaire identifies the participant’s present and past history of eye conditions and use of medications. The information will help in the assessment and diagnosis of the retinal photography findings.

MESA Retina or Retinal Photography aims at evaluating the relation of retinal microvascular characteristics (e.g., retinal arteriolar narrowing, arterio-venous nicking, and retinopathy) to subclinical cardiovascular disease, clinical disease, and their risk factors. It is propose to test new hypotheses that link retinal microvascular characteristics and arteriolar caliber to a wide array of subclinical cardiovascular measures (including left ventricular function defined from cardiac magnetic resonance imaging, peripheral arterial function defined from radial artery tonometry, and endothelial function defined from flow-mediated vasodilation of the brachial artery), clinical cardiovascular outcomes (including coronary heart disease, congestive cardiac failure and stroke), and their risk factors (including hypertension and diabetes).

II.	METHODS

1.	General Instructions

This is an interviewer-administered questionnaire. Questions should be read to the participant verbatim as they appear on the form to ensure standardization.  In addition, any introductory and transitional wording should be read verbatim.

If the topic should arise, remind the participant that all information is strictly confidential and will be used only for research purposes. Explain that information about the Eye History may be important in understanding their health. The information may, in turn, help us to better understand the causes of heart disease.

In general, for each question, possible responses are: “Yes” (and choosing either right, left or both eyes), “No”, “Don’t Know”, “Refused” or “Not Applicable”.  Have the participant choose the appropriate responses for each question. Do not probe to make interpretations about a participant’s specific symptoms. Ask questions as written and record answers as given.

2.	Specific Instructions

Begin by reading the introductory statement then proceed with the questions. Remind the participant to ask for clarification at anytime when any question is unclear.

The Eye History asks about your present and past history of eye conditions and use of eye medications.  The information will help us interpret the retinal photographs.

1.	Have you ever been told by an eye doctor that you have or had a cataract in either of your eyes?  Another term for a cataract is "opacity of the lens of the eye".

If “no”, “don’t know” or “refused” to give a response, go to question #2.

If “yes”, ask, “in your left eye, right eye, or both eyes”.

Bubble in the appropriate choice and continue with question #1a.

1a.	Did you have a cataract operation?

Cataract operation/surgery is defined as removal of the lens using any surgical techniques.

If “no”, “don’t know” or “refused” to give a response, go to question #2.

If “yes”, ask, “in your left eye, right eye, or both eyes”.

Bubble in the appropriate choice and continue with question #1b.

1b.	For the eye(s) above, when was your first cataract operation?

Choices includes: “no operation”, “don’t know”, “refused”, and “year” of surgery.

If the participant had a cataract operation, they should give the year of his/her first or only operation for each eye.

Write the year of the first operation in the blocks provided or bubble in other response in the appropriate circle and continue with question #2.

2.	For the past 3 months, or longer, have you experienced or been told you have dry eyes, where your eyes feel like something is in them, itch, burn, feel gritty, that is not related to allergies?

If “no”, “don’t know” or “refused” to give a response, go to question #3.

If “yes”, ask, “in your left eye, right eye, or both eyes”.

Bubble in the appropriate choice and continue with question #2a.

There are many manifestations of "dry eye" besides the symptoms in the question, including excess tearing. Most people will be familiar with this condition. They need not have been told by an eye doctor they have this condition.

Emphasize "for the past 3 months or longer" when you ask this question.

Emphasize "not related to allergy" when clarifying response.

If the participant has occasional episodes of dry eye, the response would be “yes” only if on average it is a complaint at least 4 out of 7 days for the past 3 months or longer.

2a.	Have you been using artificial tears for your dry eyes for the past three months or more?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

Again, note that the question requires current (4 out of 7 days) use of artificial tears for the past 3 months or longer.  In addition, there are a number of over-the-counter artificial tears which are used for this condition.

Visine, taken for allergies, does not qualify as an artificial tear unless specifically prescribed for dry eye.

Drops taken for contact lens do not qualify as artificial tears for dry eye.

3.	Has a doctor ever said you had diabetes, or high blood sugar or sugar in your urine?

If “no”, “don’t know” or “refused” to give a response, go to question #4.

If “yes”, ask, “was it confirmed as diabetes or suspected to be diabetes?”

Bubble in the appropriate choice and continue with question #3b.

This question asks whether the participant has ever been told by a health provider that he/she has diabetes, high blood sugar, or sugar in the urine. The diagnosis must be confirmed by a physician and cannot be a conjecture on the part of the participant about the conditions possibly being present.

Persons who state that they once had sugar in the urine or were told about some sugar in the urine should be asked, "Were you ever told that you have diabetes?" If diabetes is suspected, mark "suspect/possible/borderline"; if not, "no" should be marked. Gestational diabetes is not considered as being diagnosed to have diabetes and the correct response for women who give this history in the absence of a history of diabetes mellitus or "suspect diabetes" is "no".

3a.	Have you ever had laser treatment applied to the retina, the back of your eye, because of diabetic retinopathy?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye, right eye, or both eyes”.

Bubble in the appropriate choice. Continue with question #4.

“Yes”, when laser photocoagulation or laser treatment for diabetic retinopathy is specified. Xenon photocoagulation would also be coded as a "yes".

Some older persons may be confused, mistaking the bright light of either a fundus camera, slit lamp, or direct ophthalmoscope for laser treatment (these are diagnostic examinations and are not laser treatment). If you think the participant doesn't understand the question, ask him/her who did the treatment, or ask the participant to describe the treatment.

4.	Have you ever been told by an eye doctor that you have glaucoma, which is the result of high pressure in your eyes?

If “no”, “don’t know” or “refused” to give a response, go to question #5.

If “yes”, ask, “in your left eye, right eye, or both eyes”.

Bubble in the appropriate choice and continue with question #4a.

This question is specific concerning the diagnosis of glaucoma (i.e. high intraocular pressure leading to damage in the nerve fiber layer of the retina or high pressure requiring treatment with anti-glaucoma medications or surgery).

For clarification "glaucoma can be referred to as high pressure causing damage to the nerve of the eye which may be accompanied by loss of peripheral or side vision".  If a participant says "Yes, I have had some high pressure in my eyes but I have never been told that I have glaucoma", this answer would be marked as "No", because the participant may have a high intraocular pressure but not have glaucoma (intraocular hypertensive) or be an undiagnosed glaucoma patient. The diagnosis may be made by an optometrist (OD), an ophthalmologist (MD), or any other physician.

4a.	Do you take medications for your glaucoma?

If “no”, “don’t know” or “refused” to give a response, skip to question #4c.

If “yes”, bubble in the appropriate choice and continue with question #4b.

4b.	Do you use pilocarpine eye drops as part of your glaucoma medications?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye, right eye, or both eyes”.

Bubble in the appropriate choice and continue with question #4c.

Medications may be drops or oral medications usually taken to lower the pressure in the eye/s.

Pilocarpine eye drops refer to a specific type of glaucoma medications.  There are several brand names for pilocarpine drops, including Isopto Carpine, Pilostat, Pilocar and Akarpine.

4c.	Did you have surgery for your glaucoma?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye, right eye, or both eyes”.

Bubble in the appropriate choice and continue with question #5.

A glaucoma operation is defined as either a laser trabeculoplasty or trabeculectomy with the attempt to make a bleb to relieve pressure that is causing glaucomatous damage to the eye. The person is asked which eye was operated on. A peripheral iridectomy for angle closure glaucoma would also be coded as a glaucoma operation.  However, a peripheral iridectomy done as part of a cataract surgery (usually to prevent glaucomatous damage) is not.

5.	Have you ever been told by a doctor that you experienced a retinal detachment?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye only, right eye only, or both eyes”.

Bubble in the appropriate choice and continue with question #6.

A retinal detachment is defined as a separation of the retina usually due to a retinal tear or hole. Other causes, such as traction, are acceptable.

6.	Have you ever experienced or been diagnosed with age-related macular degeneration?

If “no”, “don’t know” or “refused” to give a response, skip to question #7.

If “yes”, ask, “in your left eye only, right eye only, or both eyes”.

Bubble in the appropriate choice and continue with question #6a.

This question should be asked exactly as written, using no probes. If the participant does not understand it, repeat the question, and if the participant is still not sure about the correct answer, mark "don't know".  Macular degeneration can be present without affecting vision. If subject says they have early stages of age-related macular degeneration, mark "Yes" for specific eye.

6a	Have you ever had laser treatment for macular degeneration?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye only, right eye only, or both eyes”.

Bubble in the appropriate choice and continue with question #7.

7.	Has either of your eyes been injured and required a doctor’s care?

If “no”, “don’t know” or “refused” to give a response, skip to question #8.

If “yes”, ask, “in your left eye only, right eye only, or both eyes”.

Bubble in the appropriate choice and continue with question #7a.

A “yes” response means the injury must have caused the participant to see a doctor for care.

For any trauma, such as a minor trauma or foreign body for which the participant did not see a doctor, the answer would be "no".

7a.	Was this injury from a blunt object like a fist, ball, car dashboard, etc.?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye only, right eye only, or both eyes”.

Bubble in the appropriate choice and continue with question #4b.

In ascertaining the type of trauma the participant suffered, he/she should be asked if it was a blunt trauma, which led to a visit to his/her eye doctor. A fist, a baseball, and the dashboard of a car are the given probes here. Other types of blunt trauma are also possible.

7b.	Was this injury from a sharp object like a knife, glass, or other object that cut the eye?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye only, right eye only, or both eyes”.

Bubble in the appropriate choice and continue with question #4c.

Even if the participant stated that the injury to his/her eye was from a blunt trauma, this question should still be asked. It is possible that the participant has suffered more than one type of eye injury.

Note that this question specifically asks about an injury where a sharp object penetrated the eye. For example, if an eye is hit with a fist with a sharp ring and the cornea is penetrated, 7a and 7b would both be answered "yes". If the cornea is not penetrated, "yes" for 7a and "no" for 7b. A dog bite that penetrates the cornea would be coded "yes" for 7b.

A corneal irritation or abrasion due to a contact lens requiring medical care is marked "yes" for 7 but "no" to 7a and 7b.

7c.	Was this injury due to a chemical burn, from substances like acids or lye?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye only, right eye only, or both eyes”.

Bubble in the appropriate choice and continue with question #7d.

Any chemical burn leading to seeing a doctor for care is marked as "yes."

7d.	Did this injury occur at your workplace?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye only, right eye only, or both eyes”.

Bubble in the appropriate choice and continue with question #8.

8.	How would you rate your vision without correction (without eye glasses or contact lenses)?

Responses are: “excellent”, “good”, “fair”, “poor”, “can’t see at all”, “don’t know”, “refused” to give a response.  Bubble in the appropriate choice and continue with question #9.

This question relates to the subject's perception of his/her overall vision using glasses or contact lens if the subject uses them.

9.	Do you drive at night?

If “don’t know” or “refused” to give a response, skip to question #10.

If “yes”, skip to question #9b.

If “no”, continue with question #9a.

Bubble in the appropriate choice.

9a.	Is this because of your vision?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

Bubble in the appropriate choice and skip to question #10.

9b.	How much difficulty do you have seeing things (like reading road signs) when you drive at night?

Responses are: “none”, “a little”, “a moderate amount”, “a lot”, “don’t know”, “refused”.

Bubble in the appropriate choice and continue with question #10.

10.	Have you ever been told by a doctor that you had lazy eye or amblyopia?

Responses are: “yes”, “no”, “don’t know” or “refused” to give a response.

If “yes”, ask, “in your left eye only, right eye only, or both eyes”.

Bubble in the appropriate choice and continue with question #11.

Amblyopia or "lazy eye" is usually defined as a decrease in vision in one eye, usually beginning in childhood as a result of strabismus (exotropia or esotropia) or a significant difference in the refractive (correction) error between eyes.  The subject may have had the "good eye" patched as a child to improve vision in the amblyopic eye.

- Do you have an optometrist or ophthalmologist that you go to?

If “yes”, continue with question #11a.

If “no”, “don’t know” or “refused” to give a response, bubble in the appropriate circle.  This questionnaire is now completed.

11a.	If yes, would you give his/her name and telephone number?

Record the participant’s optometrist or ophthalmologist’s name and phone number in the space provided.  This questionnaire is now complete.

A reviewer should review the form for any incompleteness.  Then complete the questionnaire form by filling in the fields in the box labeled, “For MESA Field Center Use Only”. Record the interviewer ID, reviewer ID number, and Data entry ID.

3.4.11	Neighborhood Activities

# I. Purpose

This questionnaire is designed to collect information on perceived neighborhood characteristics in order to study if features of residential environments are related to the development and prevalence of cardiovascular risk factors and cardiovascular disease. It builds on information collected in visit one, and will allow us to investigate if specific features of residential environments (such as availability of recreational spaces, or “walkability” of the neighborhood) are related to specific risk behaviors (such as physical activity). Participants are also asked about the extent to which they spend time in their neighborhoods and perform daily activities (such as food shopping) in their local areas. The responses to these questions will help us determine to what extent features of residential areas are relevant to these health behaviors.

II. Materials/Equipment

This is an interviewer-administered form. Response cards with the relevant options are available for some questions.

III. Definitions

The terms used in the questionnaire should require no explanation, since they are used in the way they tend to be used by most people in every day life.  There is no precise definition of a neighborhood. For the purpose of this questionnaire neighborhood is defined as the area within about a mile or a twenty minute walk from the person’s home. This is specified at the beginning of the pertinent questions. In some places persons may not be used to walking and may not know how far a twenty minute walk is. In this case the interviewer should emphasize that we mean about a mile from their home. When driving at 20 miles an hour, a mile is a very short drive, only about 3 minutes. In the Spanish translation, approximate equivalencies in kilometers are also given.

IV. Methods

1. General Instructions

This questionnaire is part of the Neighborhood Ancillary Study to MESA. At some sites participants will have given separate consent to participate in this study. If this is the case only participants who have agreed to take part in this study should be administered this questionnaire. At other sites, the neighborhood questionnaire has been included in the general MESA consent. In this case all participants who signed the general visit 2 (or visit 3) consent should be administered this questionnaire. Five different response cards are available. Hand the appropriate response card to the participant as detailed in the specific instructions. After handing the participant the appropriate card, read the responses on the card. If the same responses are used for several items, read the responses only the first time. A “Don’t know” option is available for several items but is not listed on the response cards and should not be read to participants.

Other important things to remember are:

- Some participants may wonder why we are asking them about their neighborhoods in a medical study. The introductory script at the beginning of the questionnaire includes some language that may help clarify this question. The paragraph below from the MESA newsletter also described why we are asking these questions.

The purpose of the MESA Neighborhood Study is to investigate if things about people’s neighborhoods are related to the presence of risk factors for cardiovascular disease or to cardiovascular disease itself. Previous studies have shown that the rate at which cardiovascular disease develops does in fact differ across neighborhoods. We do not yet know why this happens. Do things about neighborhoods cause some people to develop cardiovascular risk factors and cardiovascular disease more rapidly?  Does the availability of recreational resources in neighborhood influence the extent to which people get exercise?  Do the types of food available influence what people eat? Is living in some neighborhoods more stressful than living in others and is this related to the risk of developing heart disease? These are the types of questions that the MESA Neighborhood Study will try to answer. For this reason we will be asking you several questions about what your neighborhood is like, what things are or are not available, and what it is like to live there in general. Your answers to these questions will help us construct measurements of things about neighborhoods which may be important to cardiovascular health. Some of these questions may seem hard to answer at first, but remember that we are asking about your opinion and there are no right or wrong answers. Some may also seem repetitive, but remember that by putting together your answers to several related questions we will get a much better sense of what it is really like to live in your neighborhood. We hope that figuring out if  (and how) things about neighborhoods may affect people’s health, we will be able to improve our ability to prevent disease in the future.

- Several items may seem a bit repetitive. This is because they are part of scales. They are similar but not identical. Using several questions will allow us to measure the characteristic better than a single item. This is the reason for the redundancy.

- The terms used should be understood by most people. If the participant asks about the meaning of any item or tries to qualify a statement, please re-read the statement (or question) to them. Do not attempt to explain the question or provide synonyms (unless specified in the specific instructions below), because this may create problems for some of the scales.

- Sometimes participants may find the options hard to choose from. Don’t let them spend too much time on any one question. Ask them to choose what seems best even if it is not “exact”.

- If the participant says that he or she has only lived in the neighborhood for a short time, ask them to respond to the best of their ability. Some questions allow for “don’t know” options. Participants will also be asked how long they have lived in the neighborhood in order to take this into account when interpreting responses.

- The size of the area defined as a neighborhood in the questionnaire is relatively small. However if that participant says that things are different from one part of their neighborhood to another, ask him or her to respond with respect to most or the major part of the neighborhood.

2. Specific instructions

Begin by reading the introductory script.

The following questions refer to where you do some of your usual activities such as food shopping and exercising and what it is like to live in your neighborhood. Your answers to these questions will help us study if things about people’s neighborhoods are related to their risk of getting heart disease or stroke. This knowledge may help us develop better ways to prevent cardiovascular disease. Remember that there are no right of wrong answers to these questions. We are interested in your opinions of what it is like to live in your neighborhood.

Questions 1-9 ask about where the participant performs daily activities such as food shopping, eating out, and exercising.

1.  About how far from your home is the place (or group of places) where your household does most of its food shopping?

Hand participant the response card A and read the responses on the card.

1 mile or less

more than one mile but less than five miles

5-10 miles

more than 10 miles

don’t know

The question refers to food shopping for the participant’s household. The participant should answer in reference to their household even if he or she is not the one who does the shopping. Ask the participant to respond to the best of their ability. Only mark the “don’t know” option if they insist that they do not know.   If the participant asks what is meant by “most” of the food shopping say “the major part” or “the majority”. If they shop at several places located at different distances from their home, ask them to answer with respect to the place they use to buy the largest amount of food.

2. About how much of your household food shopping would you say is done within a 20 minute walk (or about a mile) from your home?

Hand participant the response card B, and read the responses on the card.

All or almost all of it

Most of it

About half of it

Some of it

None or almost none of it

Don’t know

If necessary clarify that when driving 20 miles an hour a mile is a very short drive (only 3 minutes or so). Ask the participant to respond to the best of their ability, even if he or she is not the one who does the shopping. Only mark the “don’t know” option if they insist that they do not know.

Note that responses are in order of decreasing frequency. Thus,  “Most of it” means less than “All or almost all of it” but more than “Half of it”. “Some of it” implies less than “Half of it” but more than  “None or almost none of it”.

3.  In what kind of food stores does your household do most of its food shopping? If you use several types of stores, please tell me the one you use most often.

Supermarkets or large superstores

Small grocery stores, Bodegas or Delis

Convenience stores (like Seven-Eleven)

Specialty stores (e.g. Italian specialty stores, health food store, or others)

Don’t know

Ask the participant to respond to the best of their ability, even if he or she is not the one who does the shopping. Do not read out the don’t know option. Only mark the “don’t know” option if they insist that they do not know.

4.  In an average week how often do you eat a meal from a fast food place  (such as McDonalds, KFC, Taco Bell, or  take-out pizza places)? By meals we mean breakfast, lunch or dinner. Include the times you take-out or order food as well as the times when you eat at the fast food place.

Hand participant response card C and read out the options.

Never ------------------------------------- skip to question 6

Less than once a week

1-2 times a week

3-4 times a week

Five or more times a week

In answering this question the participant should include take-out, order out and eat-in occasions. Chinese places and Delis are not considered fast-food places for purposes of this question.

If the response is “Never” skip to question 6.

5.  In an average week how often do you eat a meal (eat-in, order, or take out) from a fast food place which is located within a mile (or a 20 minute walk) from your home?  By meals we mean breakfast, lunch or dinner.

Hand participant response card C and read out the options.

Never ------------------------------------- skip to question 6

Less than once a week

1-2 times a week

3-4 times a week

Five or more times a week

Please note that question 5 is identical to question 4, except that it asks only about fast food places located within a 20 minute walk or about a mile from the participant’s home. Please emphasize this point.

The following questions (6- 9) refer to the participant’s exercise habits.

Now we would like to ask you some questions about your exercise habits.

6. Do you exercise at least once a week? By exercise we mean walking or bicycling for exercise, going for a jog or swim, participating in sports or exercise classes, or using training equipment or machines.

Yes___

No___    ---------------------- SKIP TO QUESTION 10

If the participant walks to work for exercise or walks to stores in order to get exercise, consider this “walking for exercise.” Please note that bicycling is also considered exercise although it is not specifically listed.

In response is “No” skip to question 10.

7. Which of the following places do you use most often to get exercise? If you use more than one type of place, please tell me the one you use most often. If you use different places in different seasons, please tell me the place you use most often overall.

__Public parks, public pools or other public recreational facilities such as community centers free and open to the public

__Streets or sidewalks

__School facilities

__Church facilities or other religious institutions

__Private gyms, clubs, dance studios

__ YMCA/YWCAs

__Your own home---------------------- SKIP TO QUESTION 10

In this question exercise is defined as noted in question 6.

If the response is “Your own home” skip to question 10.

8.  About how far from your home is the place you use most often to get exercise?

Hand participant response card A and read the options on the card.

1 mile or less

more than one mile but less than five miles

5-10 miles

more than 10 miles

Note that “the place you use most often to get exercise” refers to the one identified by the participant in question 7.

9. When you get exercise, how often do you do it within about a 20 minute walk (or about a mile) from your home?

Hand participant response card D and read the options on the card.

All or almost all the time

Most of the time

About a half  of  the time

Some of the time

None or almost none of the time

Note that this question refers to all exercise carried out by the participant (i.e. it does not refer exclusively to exercise performed in the type of  place the participant uses most often).

Please note that responses are in order of decreasing frequency. Thus,  “Most of the time” means less than “All or almost all of the time” but more than “Half of the time”. “Some of the time” implies less than “Half of the time” but more than  “None or almost none of the time”.

Question 10 asks about facilities that may or may not be available in the participant’s neighborhood in order to study if their availability is related to physical activity.

10. Now I would like to ask you some questions about facilities which may or may not be available in your area. Please tell me if there are any of the following within a 20 minute walk (about a mile) from your home.

|  | Yes | No | Don’t know |
| --- | --- | --- | --- |
| Public park |  |  |  |
| Public sports field, basketball court or tennis court |  |  |  |
| Public pool or beach |  |  |  |
| Schools, colleges, or community centers with recreational facilities that are free and open to the public |  |  |  |
| Gyms, health /fitness clubs or pools that you have to join and pay for |  |  |  |
| YMCAs or YWCAs |  |  |  |
| Bicycle path (in the street or in a park) |  |  |  |
| Are there sidewalks in your neighborhood? |  |  |  |

Please note that the item on “Schools, colleges or community centers” refers to places that are free, whereas the next item “Gyms, health/fitness clubs or pools” refers to places that one has to pay for.

If the participant states that there are sidewalks only on parts of the neighborhood, ask him or her to respond with respect to the major part of the neighborhood. For example, if most of the neighborhood does not have sidewalks, the answer should be “no”.  If most of the neighborhood does have sidewalks the answer should be “yes”.

Question 11 lists a series of statements. The participant is asked to what extent he or she agrees with the statements listed. The responses will be used to develop different indices of the participant’s residential environment. The statements should be self-explanatory. If the participant has questions on the meaning of any of the statements, attempt to clarify by repeating the statement. Do not attempt to elaborate on the statement as this may create problems in the scales. The participant should not spend too much time on any one question. The responses are ordered from strongly agree to strongly disagree.  “Neutral” or “neither agree nor disagree” is the middle option.

11. For each of the statements that I will read you now, please tell me whether you agree by choosing the best option on the card. In answering these questions, please think of your neighborhood as the area within about a 20 minute walk (or about a mile) from your home.

Hand participant response card E and read the options on the card..

|  | Strongly agree | Agree | Neutral (Neither agree nor disagree) | Disagree | Strongly disagree |
| --- | --- | --- | --- | --- | --- |
| It is pleasant to walk in my neighborhood. |  |  |  |  |  |
| My neighborhood is attractive. |  |  |  |  |  |
| There is a lot of trash and litter on the street in my neighborhood |  |  |  |  |  |
| My neighborhood is friendly |  |  |  |  |  |
| The trees in my neighborhood provide enough shade |  |  |  |  |  |
| My neighborhood has heavy traffic |  |  |  |  |  |
| There is a lot of noise in my neighborhood |  |  |  |  |  |
| In my neighborhood it is easy to walk to places |  |  |  |  |  |
| There are stores within walking distance of my home |  |  |  |  |  |
| In my neighborhood, the streets and sidewalks are in good condition |  |  |  |  |  |
| I often see other people walking in my neighborhood |  |  |  |  |  |
| I often see other people exercise (for example jog, bicycle, play sports) in my neighborhood. |  |  |  |  |  |
| I feel safe walking in my neighborhood day or night |  |  |  |  |  |
| Violence is a problem in my neighborhood |  |  |  |  |  |
| A large selection of fresh fruits and vegetables is available in my neighborhood |  |  |  |  |  |
| A large selection of low fat products is available in my neighborhood |  |  |  |  |  |
| There are many opportunities to purchase fast foods in my neighborhood (fast foods includes places like McDonalds, Taco Bell, KFC and take out pizza places) |  |  |  |  |  |

If fresh fruits and vegetables are not sold in the participant’s neighborhood the response to the item on “A large selection of fresh fruits and vegetables is available in my neighborhood” should be “strongly disagree”. If low fat products are not sold in the participant’s neighborhood the response to the item on “A large selection of low fat products is available in my neighborhood” should be “strongly disagree”.

Question 12 will be used to construct a measure of participation in local neighborhood organizations (sometimes referred to as “social capital”) which has been hypothesized to be important to health.  The list is drawn from other surveys conducted nation-wide.

12. I am going to read you a list of organizations. Please tell me if you regularly join in the activities of these organizations with people in your neighborhood.

|  | Yes | No |
| --- | --- | --- |
| A neighborhood association like a block association, a homeowner or tenant association or a crime watch group |  |  |
| Religious groups or charitable organizations |  |  |
| Parent-teacher associations or other school support or service groups |  |  |
| Youth organizations such a youth sports league or the scouts |  |  |
| Clubs or associations for senior citizens or older people |  |  |
| A labor union |  |  |
| A professional, trade, farm or business association |  |  |
| Adult sports clubs or leagues or an outdoor activity club |  |  |
| A literary, art, discussion or study group or a musical, dancing or singing group |  |  |
| Any other hobby club or society |  |  |
| Ethnic, nationality or civil rights organizations |  |  |
| Other public interest groups, political groups or party committees |  |  |

It is important to emphasize that the question refers to participation in these organizations with other people in the participant’s neighborhood. For example, if the participant is involved in a labor union at work but no one from his or her neighborhood participates, the answer to the labor union question would be “No”.

Question 13 collects information on the amount of time participants spend in their neighborhood. We need this information in order to be able to do analyses separately for persons who spend a lot of time in their neighborhood (and hence may be more affected by neighborhood conditions) and those who do not. As in the rest of this questionnaire, the neighborhood is defined as the area within about a 20 minute walk or about a mile from the participant’s home.

Time spent in the neighborhood includes time spent at home (including sleeping time) as well as time spent doing things in the neighborhood. For example, if the participant sleeps at home every night the response to the item “Monday-Friday: nights” would be “All or almost all of the time”.

If the participant says that the time spent in the neighborhood varies from week to week, ask them to respond regarding what happens most commonly.

Please note that responses are in order of decreasing frequency. Thus,  “Most of the time” means less than “All or almost all of the time” but more than “Half of the time”. “Some of the time” implies less than “Half of the time” but more than  “None or almost none of the time”.

13.  On an average week about how much time would you say that you spend in your neighborhood? When thinking of the time you spend in your neighborhood include the time you spend in  your home (including sleeping time) as well as the time you spend doing things in your neighborhood. Tell me the best option for each of the time periods I will read to you.

Hand participant response card D and read the options on the card.

|  | All or almost all of the time | Most of the time | About half of the time | Some of the time | None or almost none of the time |
| --- | --- | --- | --- | --- | --- |
| Saturday and Sunday |  |  |  |  |  |
| Day (8 am to 6 pm) |  |  |  |  |  |
| Evenings (6-9pm) |  |  |  |  |  |
| Night (9 pm-8 am) |  |  |  |  |  |
| Monday through Friday |  |  |  |  |  |
| Day (8 am to 6 pm) |  |  |  |  |  |
| Evenings (6-9pm) |  |  |  |  |  |
| Night (9 pm-8 am) |  |  |  |  |  |

Question 14 asks about the time the participant has lived in the neighborhood.

14. How long have you lived in this neighborhood? Think of your neighborhood as the area within a 20 minute walk or about a mile from your home.

##### If one year or more specify number of years (round to nearest whole year) _______

If less than one year, specify the number of months (round to nearest whole month)_________

If response is more than one year, specify the number of years. If response is less than a year, specify the number of months.

RESPONSE CARDS FOR NEIGHBORHOOD QUESTIONNAIRE

Response card A (questions 1 and 8)

1 mile or less

more than one mile but less than five miles

5-10 miles

more than 10 miles

Response card B (question 2)

All or almost all of it

Most of it

About half of it

Some of it

None or almost none of it

Response card C (questions 4 and 5)

Never

Less than once a week

1-2 times a week

3-4 times a week

Five or more times a week

Response card D (questions 9 and 13)

All or almost all the time

Most of the time

About a half of the time

Some of the time

None or almost none of the time

Response card E (question 11)

Strongly agree

Agree

Neutral (neither agree nor disagree)

Disagree

Strongly disagree

3.4.12  	Residential History

# I. Purpose

This questionnaire is designed to collect information on the participant’s residential history in order to study if exposure to air pollution over long periods is related to the development of atherosclerosis.  Information on where each participant has lived over the past 20 years or so will be linked to air pollution monitoring stations all over the country in order to measure each participant’s long term exposure to air pollution.

II. Materials/Equipment

None.

III. Definitions

None.

IV. Methods

1. General Instructions

This questionnaire is part of the Air Pollution Ancillary Study to MESA. At some sites participants will have given separate consent to participate in this study. If this is the case only participants who have agreed to take part in this study should be administered this questionnaire. At other sites, the residential history  questionnaire has been included in the general MESA consent. In this case all participants who signed the general visit 2 (or visit 3) consent should be administered this questionnaire.

Participants may be wary of providing the residential history information. If so, explain the purpose of the study and emphasize that the information they provide will be kept strictly confidential and will be used only for research. Some language from the participant newsletter explaining the study is included below.

The purpose of the MESA Air Pollution Study is to investigate if the levels of small particles present in air pollution are related to the development of atherosclerosis. Because long-term exposure to these particles may be important, we need to know where you have lived over the last two decades or so. Using this information together with information on air pollution monitoring stations all over the country, we will be able to determine if persons exposed to higher levels of particles over the past 20 years have more atherosclerosis. If this turns out to be true, it may suggest that controlling air pollution may be one more way to prevent heart disease. This is why we will be asking you for the addresses of where you have lived over the past 20 years.

Participants may not remember exact addresses, but even some information is better than none.

Take special care in the printing of numbers and street/city/state names.

Seasonal moves should be reported on separate address pages numbered sequentially.

2. Specific Instructions

Begin by reading the introductory script.

The following questions refer to the places where you have lived in the past. We are asking you these questions to study if things about the places where people have lived (such as air pollution) are related to getting heart disease or stroke. Using this information together with information on air pollution monitoring stations all over the country, we will be able to determine if persons exposed to higher levels of pollution have more disease. This information will be kept strictly confidential.

Let’s start in 1980. That would be when you were  about  ____ years old.

In order to help the participant remember, calculate approximate  age in 1980 as current age- 23 if administered in 2003 and current age-24 if administered in 2004].

Address 1

1. Where did you live in January of 1980?

Street address (Street and number)

City/Town

County

State

Zip code

Country

Complete all available information. Ask for city/town, county, state, and zip code even if participant does not remember street address. Apartment numbers are not necessary.

If the location is not in the US complete only the city/town, province, and country.

If the address is in the US but the participant does not remember the exact street address continue with the following question:

We realize it may not be easy to remember the exact street address. Do you remember the nearest intersection?

_____________________________________  and __________________________________

Please be sure to include the suffix (for example Street, Avenue, or Boulevard) together with the street name.

Questions 2-4 collect information on smoking. In studying the possible effects of air pollution, it is important to know if the participant (or people who lived or worked with him or her) smoked. The questions refer to the time when the participant lived at the address provided above.

2. Did you smoke at the time you lived at this address? Yes____    No_____

If necessary clarify by saying “If you quit or started smoking while you lived there tell me if you smoked most of the time you were there”.

3. Did anyone else living with you smoke inside your home while you were living at this address?

Yes____  No____

If necessary clarify by saying  “If this changed while you lived at this address, tell me if anyone smoked for most of the time you lived at this address. ”

4. If you were working, did people who worked with you smoke at work?

Yes____ No____ Did not work____

If this changed over the period, ask about the majority of the time.

5.  Did you live anywhere else after you lived at this address?

Remind the participant what address you are referring to by reading off the last address reported.

NO------------------------------ Confirm that this is the participant’s current address. If it is, 					QUESTIONNAIRE ENDS HERE

Thank you so much for completing this questionnaire!

This information will be kept strictly confidential and will 					help us	understand if  air pollution is related to heart disease 					and stroke.

If the response is “NO”, and this is the participant’s current address, the questionnaire is over.

If this is NOT the participant’s current address, it means that interim addresses (between the last address on the form and the participant’s current addresses) may have been missed. Probe for additional addresses after the last address reported, change response to question 5 to YES, and complete additional address pages as necessary. THE LAST ADDRESS REPORTED SHOULD BE THE PARTICIPANT’S CURRENT ADDRESS.

YES---- CONTINUE TO NEXT QUESTION

6. When did you move from this address?           Month___________  Year___________

Complete the month and year in which the participant moved from the address recorded above.

The cycle of questions 1-6 is repeated for the next address until the current address is reached. Successive addresses are numbered consecutively: Address 1, Address 2, Address 3 etc.

Situations when the participant reports living at two addresses over a given period:

Seasonal moves:

Separate forms should be filled out for seasonal moves (i.e. each address is captured on a different form), as long as the participant spends 3 months or more at each address. For example, if the participant lives in Chicago but spends one month each winter in Florida, there is no need to capture the Florida address. If the participant spends 3 months each year in Florida, those addresses should be captured on separate forms, noting the time of each move. The addresses will be numbered sequentially with no overlap, although addresses may be repeated. (If capturing this information becomes too tedious during the interview, the interviewer can make a note of the two addresses and the months of the year usually spent at each and complete the individual address forms later).

In the special case when the seasonal moves are ongoing and extend back in the past the following procedure can be followed to capture both addresses without completing many forms.

Suppose the participant moved to an address in Chicago in January 1990 and has been spending the summer months (May-September) in Florida every year since then:

- Complete one supplementary address form for the Chicago address with the corresponding sequential number (based on prior addresses reported). Complete questions 2-4 for the time the participant has spent at this address. Since this is an ongoing living arrangement (i.e. it is current) the answer to number 5 will be “No”.  On question 6 specify the month the participant moves from Chicago to Florida (in this case May, or 05 for the month). Leave the year field blank.

- Complete a second supplementary address form with the same sequence number for the Florida address. Complete questions 2-4 for the time the participant has spent at this address. Since this is an ongoing living arrangement (i.e. it is current) the answer to number 5 will be “No”.  On question 6 specify the month the participant moves from Florida to Chicago (in this case September, or 09 for the month). Leave the year field blank.

Please note that this option for capturing seasonal moves applies only when the participant moves back and forth between the same addresses and this is an ongoing (current) situation. If this is not the case, it will be necessary to fill out one form for each address.

If the moves are not seasonal, but the participant reports living at two places for a given period of time:

1. Ask them where they lived most of the time and obtain that address.

2. If they say they split their time equally between both places (e.g. go back and forth very often between both, for example part of the week in one place and part of the week in another), ask if the zip code was the same:

If the zip code is the same for both, choose one address and complete that one, ignoring the other one.

If the zip codes are not the same (or the person is not sure about the zip code) fill out one address form for each address with the SAME sequence number at the top. Question 4, 5 and 6 will be identical for both address forms.

3.5  	Clinic Examinations

3.5.1 	Anthropometry

I.	Purpose

Anthropometry was obtained in the MESA “Baseline Exam” or Exam 1.  The purpose is the same as for Exam 1 and is obtained in MESA exam 2 for longitudinal studies.

II.	Materials and equipment

	Stadiometer (Accu-Hite Measure Device with level bubble) (height ruler with triangle level is used at some centers)

	 Detecto Platform Balance Scale in lbs/kg

	Gulick II 150 cm anthropometric tape

	Full length mirror

	Four 50-pound weights (certified prior to first MESA visit) to calibrate scale

III.	methods

Methods for completing the Anthropometry portion of Exam 3 are the same as in Exam 2.  Some important points are reiterated here.

General Instructions:

For all measurements, participants should wear light clothing but no shoes (thin socks or “pillow slippers” OK). Have participants completely empty their pockets and remove excessive amounts of jewelry that could affect the weight measurement. Provide lockers with locks for valuables.

Pregnant women should not be measured, regardless of gestational stage (check exclusion criteria for pregnancy). The Clinic Coordinator should ascertain pregnancy status, both for measurements and for subsequent coronary calcification measurement.

Specific Instructions:

1.	Standing Body Height ~ procedure is the same as for Exam 1 & 2.

1.1	Equipment

	Stadiometer (Accu-Hite Measure Device with level bubble) (height ruler with triangle level used at some centers is adequate)

1.2	Before measuring height, make sure the floor is level, the wall is at a 90 degree angle to the floor, the wall is straight, and the Stadiometer is mounted perpendicular to the floor.

1.3	For accurate measurement of height, the participant must be standing in a vertical plane.  Please refer to the Baseline Exam manual for details. Record the results, to the nearest tenth (0.1) of a cm, in Box 1a on the Anthropometry Form.

1.4	If any modification was made to obtain height, bubble in “yes” to the question, “Was there a modification in protocol?”

2.	Body Weight ~ procedure is the same as for Exam 1.

2.1	Equipment

	Detecto Platform Balance Scale in lbs/kg

2.2	Always balance the scale so that the indicator is at zero when no weight is on the scale. The scale should be on a firm, level surface. Instruct the participant to stand in the middle of the platform of the balance scale, with head erect and eyes looking straight ahead. Adjust the weight on the indicator until it is balanced. Record the results, to the nearest 0.5lbs, in Box 2a.

2.3	If any modification were made to obtain weight, bubble in “yes” to the question, “Was there a modification in protocol?”

2.4	For detailed instruction or questions, please see the “Baseline Exam MOP”.

3.	Girth Measurements ~ procedure is the same as for Exam 1.

3.1	Equipment

	Gulick II 150 cm anthropometric tape

	Full length mirror

3.2	Technique

- Do not take abdominal and hip girth measurements over loose clothing.  It is ok if taken over light well-fitted clothes.

3.3	Abdominal Girth

Apply a Gulick II anthropometric tape horizontally at the level of the umbilicus and instruct the participant to breathe normally. Move to the participant’s right side to take the measurement; do not take this measurement from the front. Be sure to keep the tape horizontal while making the measurement; use the wall-mounted mirror to assure horizontal placement on all sides.

Round abdominal girth measurement to the nearest 0.1cm and record in Box 3a.

If the circumference exceeds 150 cm, record “yes” for the question, “Was there a modification in protocol?”

3.4	Hip Girth

Take the hip girth measurement from the participant’s right side; do not take this measurement from the front. Instruct the participant to stand with his/her feet together. Measure hip girth at the maximum circumference of the buttocks. Check to see that the tape is level in front and back.

Round hip girth measurement to the nearest 0.1cm and record in Box 3b.

If the circumference exceeds 150 cm, record “yes” for the question, “Was there a modification in protocol?”

4.	Comments/Modifications to the Protocol

If you have comments or if there have been modifications to the protocol as described above, answer “yes” to question 4 on the Anthropometry Form and record comments in the space provided. If there are no comments or modifications, answer “no” to question 4.

5.	Completing the “For MESA Field Center Use Only” section

Make sure to record the Technician ID#, Reviewer ID#, and Data Entry ID# is these fields at the bottom of the form.

6.	Quality Control ~ Calibration Check of Scales and Equipment Check

6.1	Equipment:

	Four 50-pound weights (certified prior to first MESA visit) to calibrate scale

	Gulick II anthropometric tapes

  - Check scales for accuracy on a monthly basis.

6.21	Place two weights on the scale and record the numeric value obtained in the “Light Poise” column of the “Scale Calibration Checklist.” Add two more weights and record the numeric value obtained in the “Heavy Poise” column.

6.22	The values obtained should be within 1.0 pound of the expected weight. If either value exceeds this limit, the scale must be calibrated by the manufacturer or by the appropriate institution personnel.

6.23	When the scale is not in use, keep it balanced at 300 pounds. This keeps the tension off the internal spring mechanism.

  - Examine anthropometry tape measures on a weekly basis for sign of wear.

3.5.2  	Seated Blood Pressure

I.	Purpose

Seated Blood Pressure was obtained in the MESA “Baseline Exam” or Exam 1.  The purpose is as for Exam 1 and is obtained in MESA Exam 3 for longitudinal studies.

Again, the Dinamap automated device will be used for consistency and to reduce the potential for observer biases.

II.	Materials and equipment

	Dinamap automated blood pressure device (Dinamap Monitor Pro 100, which includes printer paper, power cable, and power converter.)

	Blood pressure cuffs in a variety of sizes (Dura-cuf Adult Assortment Pack [#2699]).

	Measuring tape (for arm circumference).

	Watch or stop watch (to time five-minute rest and resting heart rate).

	Hand calculator (to average 2nd and 3rd BP readings).

	Copy of Critikon chart for choosing correct BP cuff size (see Table 2).

	Information sheet on interpretation of BP from JNC VI (see Table 1).

	Resting Heart Rate/Blood Pressure Form.

III.	Definitions

1.	Sphygmomanometry: Measurement of blood pressure.

2.	Oscillometric device: Method for measuring blood pressure that relies on the oscillation or fluctuation in arterial pressure generated by the cardiac cycle and transmitted to an inflated blood pressure cuff overlying an artery. This method differs from the auscultatory method, which relies on audible changes over an artery during deflation of an inflated cuff.

IV.	Classification of the Participant's Blood Pressure within the JNC VI Categories and criteria for alerts and referrals

This classification and the criteria for alerts have not changed from Exam 1.  However, they are important and are reiterated here.

The 1997 Report of the Joint National Committee on Detection, Evaluation, and Treatment of High Blood Pressure (JNC VI) defines categories of blood pressure and recommends follow-up according to the following criteria:

Table 1. Classification of BP in Adults Aged 18 Years or Older*.

| BP Category | SBP (mm Hg) |  | DPB (mm Hg) | Action |
| --- | --- | --- | --- | --- |
| Optimal | <120 | and | <80 | Recheck in 2 years |
| Normal | <130 | and | <85 | Recheck in 2 years |
| High-normal | 130–139 | or | 85–89 | Recheck in 1 year |
| Hypertension** |  |  |  |  |
| Stage 1 | 140–159 | or | 90–99 | Refer within 2 months |
| Stage 2 | 160–179 | or | 100–109 | Refer within 1 month |
| Stage 3 | >180 | or | >110 | Refer within 1 week or immediately |

* 	When recommendation for follow-up of DBP and SBP are different, the shorter recommended time for recheck and referral should take precedence. This classification applies only to participants not taking antihypertensive drugs.

** Diagnosis of hypertension must be based on two or more readings taken at each of two or more visits following an initial screening.

SBP= systolic blood pressure.  DBP= diastolic blood pressure.

1.	Alert levels requiring immediate referral (send participant directly to a physician or hospital) for MESA participants are:

	Systolic BP >210 mm Hg

	Diastolic BP >120 mm Hg

2.	Alert levels requiring urgent referral (within one week) are:

	Systolic BP 180–210 mm Hg

	Diastolic BP 110–120 mm Hg

3.	Alert levels requiring follow-up within two months time, and, therefore, we recommend physician notification for systolic or diastolic BP above these levels.

- BP >140/90 mm Hg

4.	JNC VI states that blood pressure classifications and referral recommendations are based on the average of two or more readings on two or more occasions. In MESA we intend to use the average of the 2nd and 3rd blood pressure readings (see below) in order to reduce the impact of reactivity (higher first reading) on the estimate of the value of the underlying blood pressure. Thus, in deciding whether a participant meets criteria for an alert level, the average of the 2nd and 3rd readings should be used. This will require on-the-spot arithmetical manipulation of the systolic and diastolic values. A hand calculator may be useful. The data forms include fields for these averaged values and for any actions taken.

V.	Methods

1.	Preparation

1.1	Record the date of the procedure and the Dinamap number on the Seated Blood Pressure Form during the five-minute rest period.

1.2	Before the BP measurement procedure, explain to the participant what to expect and how long the procedure will take. The following script is suggested:

This part of the exam involves taking your resting blood pressure. It will take about 10 minutes. We would like you to sit with both feet on the floor and your arm supported on the table. We will have you sit quietly for five minutes. Then we will take your blood pressure three times, one minute apart, using an automated device. We will give you your blood pressure readings and some material to help you interpret them at the end.

  - Make sure the room temperature is between 70 and 76 Fahrenheit.

2.	Cuff Size Selection

2.1	Use the proper cuff size to avoid under- or over-estimation of the correct blood pressure. Selection of the proper sized cuff is based on the guideline that the length of the inflatable bladder in the cuff should be at least 40% of the arm circumference. Measurement of the bladder length in the Critikon cuffs confirms that the chart in Table 3 conforms to this guideline. A copy of this chart should be available during the BP measurement procedure for easy reference. Selection of cuff size should be based on the Critikon chart in Table 2, and only Critikon cuffs should be used. If the participant’s arm size falls in a range in which there is overlap of two Critikon cuff sizes, use the larger cuff.

2.2	Measure the right arm circumference as follows:

	Ask the participant to bare the upper arm.

	Ask the participant to sit or stand holding forearm horizontal, i.e., parallel to the floor.

	Measure arm length from the acromion (bony extremity of the shoulder girdle) to the olecranon (tip of the elbow) using a metric tape.

	Mark the midpoint on the dorsal (back) surface of the arm.

	Ask participant to relax arm along side of the body.

	Draw the measuring tape snugly around the arm at the midpoint mark, keeping the tape horizontal. Only pull the tape snug enough so that the first red-bead marker can be seen.  Tape should not indent the skin. If you can see both bead, the tape is too tight. Record the arm circumference measured to the closest (0.1) cm in Field 1 on the Seated Blood Pressure Form.

	Use the criteria in Table 2, below, to determine cuff size. Check the cuff size used in Field 2 on the Blood Pressure Form by filling in the appropriate circle.

Table 2. Cuff Size Indicated by Measured Arm Circumference

| Arm Circumference* (cm)* | Cuff Name** | Bladder Length (cm) |
| --- | --- | --- |
| 12-19 | Child | 8 |
| 19.1-25 | Small Adult | 10 |
| 25.1-33 | Adult | 13 |
| 33.1-40 | Large Adult | 17 |
| 40.1-50 | Thigh |  |

* These circumferences are printed on the corresponding cuff for verification.

** Critikon Dura-cuf nomenclature is also printed on the cuff.

3.	Setting up the Dinamap BP Machine

3.1	Refer to the “Baseline Exam MOP” for further details.

4.	Positioning the Participant

4.1	The workstation should be free of excessive noise or distractions.

4.2	The participant should be seated and relaxed in a comfortable chair, to ensure that:

	He or she is sitting up (not slouched).

	Both feet are on the floor (legs/ankles not crossed).

	Right forearm is supported resting on the table.

4.3	The participant should not talk, eat, or drink during the procedure.

4.4	Ideally, the Dinamap output will not be visible to the participant during the measurement, as this may cause anxiety.

5.	Application of the Blood Pressure Cuff

5.1	Procedure unchanged from Baseline Exam.  Refer to the “Baseline Exam MOP” for further details.

6.	Rest Period

6.1	The participant should rest for five minutes (timed using a watch or stop watch) prior to the heart rate and blood pressure measurement.

6.2	When the five-minute rest period is over, but before the first blood pressure measurement is started, record the time of day on the Seated Blood Pressure Form (examples: 04:25 P [p.m.] or 11:38 A [a.m.]).

6.3	Record the room temperature on the Seated Blood Pressure Form.

7.	 Blood Pressure Measurement

7.1	Procedure unchanged from Baseline Exam.  Refer to the “MESA FC MOP Exam 1” for further details.

8.	BP Measurement Instructions for Participants With Short, Thick Arms

8.1	Occasionally there will be a participant whose upper arm is too thick and short for the thigh cuff or on whom the thigh cuff pops open on inflation. The alternative procedure in this case is to obtain the resting blood pressure in the right forearm.

8.2	Measure the forearm circumference at the midpoint between the olecranon (elbow) and the ulnar stylus (wrist bone on pinkie side). Select the proper size cuff based on the forearm measurement. The blood pressure procedure is otherwise the same.

8.3	You must document on the Seated Blood Pressure Form that you have measured the forearm blood pressure.

9.	Reporting Blood Pressure Results to Participants

9.1	The technician may verbally provide the participant with the blood pressure reading (the average of the last two pressures), if asked, after the procedure has been completed.

9.2	Alternatively, if the blood pressure is normal (<140/90), the technician may say that it is normal, particularly if asked.

9.3	If the blood pressure is not normal (>140/90) but not at an alert level (>210 mm Hg), the technician should exercise the standard option of not discussing the interpretation or stating that it does appear to be high (or “somewhat elevated”) but that, again, it will be discussed later.

9.4	If an alert level is identified, the technician should calmly notify the clinic coordinator when the procedure has been completed. (If symptoms of severe hypertension are present, the technician should notify the clinic coordinator immediately.)

VI.	Quality Assurance/Quality Control Procedures for Dinamap Pro 100

1.	Once a week each device should be used simultaneously with a paired device to simultaneously measure the blood pressure in each arm of a non-smoker under the age of 50, in whom there is no reason to suspect that the blood pressure in the two arms should differ. Repeat the measurement three times.

2.	Procedure unchanged from Baseline Exam.  Refer to the “Baseline Exam MOP” for further details.

3.5.3  Ankle/Arm Blood Pressure

I.	Background and Rationale

The presence of peripheral vascular disease will be assessed with the ankle/arm blood pressure index (ABI). A low ABI is highly correlated with lower extremity arterial disease and has been shown to predict all-cause and cardiovascular mortality in several studies. A normal ABI is 1.0 to 1.5, with progressively lower values corresponding to worsening arterial disease. This method may underestimate, however, the occurrence of the disease in patients with long-standing diabetes, due to medial arterial calcification and incompressibility of blood vessels.

II.	Equipment and Supplies

	Nicolet Doppler apparatus (EN50 LE 100, Nicolet vascular, Golden, CO) with an earpiece.

	gV batteries.

	Two full tubes of ultrasound transmission gel. (Gel can be purchased in bulk and tubes refilled.)

	A mercury sphygmomanometer with a male quick release coupler. The mercury column manometer should be mounted at “eye level.”

	Blood pressure cuffs: adult, large adult, and thigh

	Arm cuffs: Two 16 cm wide arm blood pressure cuffs (large arm), two 12 cm wide arm blood pressure cuffs (regular arm or adult), two 10 cm wide arm blood pressure cuffs (pediatric or child), and two thigh blood pressure cuffs.

	Tissue or wash cloth to remove the ultrasound contact gel.

	Black ball point pen and a marker.

	Ankle-Arm Blood Pressure Form.

III.	definitions

1.	Peripheral vascular disease, peripheral atherosclerosis, and peripheral arterial obstructive disease are synonyms. Peripheral vascular disease does not refer to venous disease, small-artery obstructive disease, vasospastic disease, cold sensitivity, or capillary disease.

2.	The ABI is a ratio of ankle to arm pressure and is computed separately for each leg. The numerator for the right leg is the higher of the two systolic ankle pressures (posterior tibial or dorsalis pedis) in the right leg. The numerator for the left leg is the higher of the two systolic ankle pressures in the left leg. The denominator for both legs is the average of the right and left brachial systolic blood pressures. If the two arm pressures differ by 10 mm Hg or more, take the higher arm pressure instead as the denominator.

3.	An ABI of 0.90 or less is considered peripheral arterial disease, although peripheral arterial disease could exist if the ABI is 0.95 or less.

4.	Moderate to severe lower extremity arterial obstructive disease is considered to be present if the ABI is less than 0.80.

5.	Medial calcification could exist in an artery, if the ankle systolic blood pressure exceeds 290 mm Hg in any participant or 240 mm Hg in a participant with brachial systolic pressure less than 160 mm Hg; or if the ABI index is >1.5.

IV.	Methods

1.	Preparation

1.1	Thoroughly explain the procedure to the participant and allow him/her to ask questions.

1.2	Conduct the examination in a quiet, warm, and comfortable room. If the room is cool, a blanket may be used to cover the participant (including arms, hands, and feet), except while the actual measurement is being made. Have the participant lie supine, with arms and legs (to mid-calf) bared, on a comfortable horizontal examination table.

1.3	Do not place blood pressure cuffs over any open lesion that could be a potential source of contamination.

1.4	Have the participant rest quietly for at least 5 minutes before beginning the measurement procedure. Record the date of the examination on the Ankle-Arm Blood Pressure Form. Before you begin the procedure, instruct the participant to remain relaxed and to refrain from helping you (e.g., lifting the arm to facilitate placement of the cuff). Once you begin the procedure, explain the steps to the participant as you proceed.

1.5	Place an appropriate blood pressure cuff around the arm, based on arm circumference at midpoint:

	Adult (12 cm width) for arm circumference of >32 cm

	Large adult (17 cm width) for arm circumference of 32–42 cm

	Thigh (20 cm width) for are circumference of 43 cm

For the ankle, use the adult (12 cm) cuff size. Place the cuff so that the lower portion rests 3 cm above the medial malleolus (ankle bone).

2.	Arterial Blood Pressure Measurement

2.1	These first two steps are optional: Locate the brachial artery on both arms by palpation in the antecubital fossa.  Similarly, locate the dorsalis pedis (dorsum of the foot) and posterior tibial (medial ankle) arteries on both legs. Mark the location of each artery with a black marker. Sometimes the arterial location in the feet will not be palpable but can be determined with the Doppler.

2.2	Using the procedure below, measure systolic arterial pressures in the following order:

	right brachial artery

	right dorsalis pedis

	right posterior tibial

	left dorsalis pedis

	left posterior tibial

	left brachial

2.21	Place a liberal amount of ultrasound conducting gel over pulse location.

2.22	Place the Doppler probe over the artery and then turn on the Doppler. Angle the probe in line with the artery and move it from side to side until the strongest pulse is heard. Don’t press too hard on the artery with the probe. Rest your hand comfortably so that the probe is secured in place once a strong pulse is heard.

	In a small percentage (< 10%) of participants, you will not be able to find the dorsalis pedis pulse. If you are having trouble, be patient and continue to search for at least three minutes. If you are still unable to locate a pulse here, enter “999” in Field 1 on the Ankle-Arm Blood Pressure Form.

2.23	Inflate the cuff slowly until the pulse is no longer audible. Inflate to 20 mm Hg  above the level at which pulse sound disappeared. (If the pulse cannot be obliterated, you may raise pressure to a maximum 300 MMHg. If not obliterated at that point, record “unable to occlude.”)  Deflate the cuff slowly allowing the pressure to drop at a rate of 2 mm Hg per second. Record the pressure at which the first sustained (more than one beat) pulse reappears. This is the systolic pressure at this location. Deflate the cuff completely. Record the measurement in Field 1 on the Ankle-Arm Blood Pressure Form. Wait for 20 seconds and then repeat the process to obtain a pressure measurement at each of the remaining sites.

2.24	If the signal remains faint as more pressure is released or if the probe moves off the artery, deflate the cuff completely, wait for 20 seconds, and then repeat the measurement.

2.3	Record the outcome (completed/not complete) in Item 2. If  not completed, continue to Item 3 and record the reason(s) that the procedure was not completed.

3.5.4  	Phlebotomy & Spot Urine Collection (see Laboratory, section 3.7)

3.6	MESA-Eye

For Exam 3, MESA-Eye exams (Retina and Refraction) will be performed on only the MESA participants (probands) who have relatives in the MESA Family Study.   Please see the MESA Family Study Exam Manual for instructions and protocols with these eye exams.

3.7  	Laboratory

Most of the procedures and laboratory equipment and supplies used in Exam 3 are similar to those in Exam 2.  Only some important components or those that have changed are re-iterated here.  Please refer to the Exam 1 and Exam 2 MOP for a detailed list.  In addition, any component new in Exam 3 will also be described.

I.	purpose

MESA is a multicenter, longitudinal epidemiological study of the incidence and progression of subclinical atherosclerotic cardiovascular disease.   The Central Blood Analysis Laboratory (CBAL) will have responsibilities for special blood collection and handling protocols as well as training and QC monitoring at the Clinical Centers.  The laboratory will also be responsible for performing assays and reporting results.

The blood samples collected and processed by Clinical Center technicians are the foundation for all of these tests.  The most important step (and potentially the most variable) is the collection and processing of the blood samples.   If the blood sample itself is not correctly drawn and processed, the laboratory results may not be precise or may not be valid.

MESA involves the collection of 49.5 mls of blood from participants at Exam 3.

II.	Equipment & Supplies

The following supplies will be provided in bulk by the CBAL:

5 ml SCAT-I tubes (1 per participant).  Must be stored refrigerated until used.

Cryovials – 0.5 ml and 1.5 ml with color-coded caps

Tubes – 4.0 mL for urine

The blood collection area should have the following supplies:

Lab coats and gloves

Phlebotomy chair

Basin (just in case)

Washcloths/Towels

Smelling salts

Lab mats and wipes

10% bleach solution or approved biohazard disinfectant

Plastic cart with wheels for phlebotomy supplies (or plastic tray with compartments)

Butterfly needles (21 G) with luer adapter  (B-D # 7251)

Vacutainer barrels

Tourniquets

Alcohol prep pads

Gauze (2x2)

Surgical tape - paper tape (easier on participants)

Band-Aids

Blood collection tubes (keep extras on hand):

2 – 10 ml Serum tubes (B-D 366430)

2 – 10 ml EDTA tubes (B-D 366457)

1 – 4.5 ml Citrate tubes (B-D 366415)

1 – 5 ml SCAT-I*

For Blind Duplicate Samples:	5 ml Serum (B-D 366534),

5 ml EDTA (B-D 366452),

4.5 ml Citrate (B-D# 366415),

5 ml SCAT-I*

Blood tube rocker

Blood tube racks

Ice bucket and crushed ice - filled 10 min before draw

Stopwatches or timers (ex. Fisher Scientific # 06-662-9)

Scissors

Pens

Labels

Phlebotomy / Processing Form

Blood Spill Kit

Biohazardous waste container

Needle/sharps container

Material for urine collection

* provided by CBAL

III.	methods

1.	Safety Issues and Precautions for Handling Blood Specimens.

In accordance with the OSHA regulations on bloodborne pathogens, the CBAL recommends the following laboratory safety protocol for the field center laboratories:

Use of non-permeable lab coats, latex gloves, and face shields when handling any blood in any situation where splashes, spray, spatter, or droplets of blood may be generated and eye, nose, or  mouth contamination can be reasonably anticipated.

Use of aerosol containers in all centrifuges.

Follow 'Universal Precautions' when handling any blood products.

Contaminated needles and sharps shall be immediately placed in a puncture-resistant, leakproof container.  Never recap or break needles.

Hepatitis B vaccine be offered to all unvaccinated technicians handling blood, and documentation of vaccination, or technician’s declining to be vaccinated, should be kept on file at the Clinical Center.

2.	Participant ID Labels

The Coordinating Center will supply each field center with sheets of sample ID barcode labels to use for labeling draw tubes, working tubes, cryovials, and freezer boxes. There will be a total of 34 labels: 6 labels for the draw tubes plus 3 extra labels, 2 labels for two pooling tubes, and 23 cryovial labels. Each participant set of barcode labels has the same 7-digit sample ID number (the first digit identifies the clinic – Wake Forest =3,

Columbia = 4, John’s Hopkins = 5, UMinn = 6, Northwestern = 7, UCLA = 8).  The cryovial labels also have a 2-digit extension (01 to 23) that serves as a unique identifier for each cryovial within a sample ID – this helps in tracking the repository.  See Appendix for proper orientation of the barcode label on the cryovial.   The extra labels are used for the urine collection cups.

There will also be special QC ID labels for the blind duplicate samples.  See the section on blind duplicates for further information on the procedure.

It is essential that blood and urine samples be precisely labeled throughout the collection and processing stages to ensure that participant samples are not miscoded. To facilitate accurate labeling, it is suggested that you pre-label sets of collection tubes and cryovials prior to the participant's visit, with a crosscheck of the labels with each participant’s ID # prior to the phlebotomy.

3.	Forms

The purpose of the Phlebotomy/Processing Forms (P/P Forms) is to facilitate the efficient collection of plasma and serum samples from the participants, with maximum protection for the participant and the technician. In addition, the P/P Forms facilitate the monitoring of phlebotomy and other quality assurance parameters and provide information critical to the interpretation of the assay results and maintenance of the sample repository.

The completed Processing Forms will be included in the sample shipments to CBAL.

Both the Phlebotomy Form and the Processing Form must be labeled with the participant ID #.  All forms must be completed in ink.

The Phlebotomy Completion Form will be scanned, and the information will be electronically sent to the Coordinating Center and the CBAL. The completed Phlebotomy Completion Form and MESA Processing Form will then be sent with the sample shipments to CBAL. Both forms must be labeled with the correct pre-printed barcode sample ID label. All forms must be completed in ink.

4.	Participant Refusal of Phlebotomy or Urine collection

Rarely, a participant will refuse to provide a urine sample or will refuse phlebotomy.  Please keep a list of MESA Enrollment ID #s of any of these participants and identify which test they refused.

5.	Urine collection & Venipuncture

5.1 	Initial preparation for specimen collection prior to the arrival of participants is similar to that of Exam 1 and 2.

5.2	Urine Collection & Processing

# 5.2.1	Preparation of Participants for Urine Collection

Urine should be collected before venipuncture; preferably as early in the visit as possible.  Urine will be collected on all participants.  Collection will be random and therefore considered a ‘spot’ urine collection.  Samples should not be collected after exertion or an acute fluid load (>24 oz), however participants should be encouraged to stay hydrated (water only) even while fasting for the visit.  Participants having difficulty producing a urine specimen may be offered a glass of water, and a second (and third) urine specimen may be collected later in the visit to bring the volume up to the required amount.

It is suggested that female participants may prefer to use the Sage Commode Specimen & Measuring System #2500 for urine collection.  If so, follow the instructions provided with this collection kit. Male participants, and female participants if they would prefer to, may urinate directly into a specimen collection container.

Do not collect urine from females who are menstruating. Collect a sample at a later visit, if possible.

Specimen containers for routine random urine collection should be chemically clean, and have a tight fitting lid to prevent leakage during transportation.  (Suggestion for standard sterile specimen container:  Lab Safety Cat# OR-29091).  The urine container must be clearly labeled with participant ID#.

    - Urine Collection:

Instructions for Participants:

The participant’s privacy should be assured.

The following steps should be followed:

- Wash hands before and after voiding.

- Preferably cleanse area with antiseptic wipes prior to voiding (optional).

- Remove the cap from the collection container and void into the container.

- Carefully seal the cap of the container so that it is tight and leak proof.

- Bring filled container to MESA staff for processing.

##### 5.2.3	Urine Processing:

## Keep urine refrigerated or on ice until processing.  One cryovial of urine is aliquoted for microalbuminuria testing using 3.0 mL urine into a large, 4.0 mL tube.   Freeze immediately after aliquoting.  Discard any extra urine.

NOTE:  Do not OVERFILL the urine aliquot since the maximum capacity is 4.0 ml. There needs to be adequate space for the urine to expand when frozen without displacing the tube cap.

5.3	Priority of tubes & Preparation of phlebotomy draw-tubes and aliquot racks

A total of approximately 49.5 ml of blood will be drawn from each participant in 6 tubes.  (20% will have 7 tubes collected for a total of 54.5 ml of blood)

The order in which the tubes are collected is important.  Blood collection must be drawn in the following order:

- 10 ml EDTA				purple top

- 10 ml Serum				red top

- 4.5ml Citrate				blue top

- 10 ml EDTA				purple top

- 10 ml Serum				red top

- 5 ml SCAT-I				yellow top

Note:  The numbering and the priority of tubes has changed from Exam 2 to Exam 3.  In addition, SCAT-1 now has a yellow top (rather than red, as it was in Exam 2).

Tubes #1 and #4 are 10ml EDTA tubes (B-D#366457).  After centrifugation, plasma from these tubes will be pooled and aliquoted into seven purple-coded aliquots.  This plasma will be used for lipid levels and banked for future testing. For selected Blind Duplicate participants, an additional 5 ml EDTA tube (BD# 366452) will be used.

Tubes #2 and #5 are 10ml red-topped Serum tubes (B-D# 366430). After filling, let these tubes stand at room temperature for a minimum of 40 minutes to allow the blood to clot. They are then centrifuged and the serum is pooled and aliquoted into seven red-coded aliquots. The serum will be tested for glucose and stored in repository for later analysis. For selected Blind Duplicate participants, an additional 5 ml Serum tube (BD# 366534) will be used.

Tube #3 is a 4.5ml blue-topped citrate tube, a silicon-coated glass tube containing 0.5ml of 3.2% sodium citrate (B-D# 366415).  After centrifugation, plasma is aliquoted into 4 blue-coded aliquots.  The plasma will be banked for future testing.

Tube #6 is a 5ml yellow-topped "Special Coagulation" tube (SCAT-I) containing a white, powdered anticoagulant provided by CBAL.  This tube contains a special combination of anticoagulants that ensure long-term stability of the plasma sample.  Specifically, this tube, when filled, will contain 4.5 mm EDTA, 150 KIU/ml aprotinin and 20 uM D-Phe-Pro-Arg-chloroketone.  The SCAT-1 tube must be stored refrigerated. It must be drawn after at least one other tube has been drawn. Important to note, this tube is 'non-sterile', therefore it must be drawn using a butterfly apparatus with 12 inches of tubing; alternatively, a syringe may be used for the venipuncture, and expressed through the cap (with great care to limit turbulence) into the SCAT tube.   It is critical that the SCAT-1 tube is mixed well (>30 sec of gentle inversion) before being placed on ice to await further processing/centrifugation.  The four aliquots for this tube are yellow-coded, and the plasma will be stored in repository for future testing.

Summary of Blood Mixing During Venipuncture

Each tube should be treated as follows:

Tube #1and 4	EDTA – place on mixer for ~30 seconds, then place in ice bath.

Tube #3	Citrate – place on mixer for ~30 seconds, then place in ice bath

Tube #6	SCAT-1 – place on mixer for AT LEAST 30 seconds, then place in ice bath.

Tube #2 and 5	Serum – do NOT mix; place in rack at room temperature for AT LEAST 40 minutes

5.4	Collection of Blind Duplicate Tube

As with Exam 1, 20% of participants will have an additional tube of blood collected, for a total of nine tubes (approximately 55 mL) of blood. This sample is collected for quality control purposes. This sample is collected last into the Blind Duplicate Tube (#7), which may be purple, red, yellow, or blue.

5.5	Preparation of Phlebotomy Room – is similar to that of Exam  1 and 2.

Setup of Draw Tube and Aliquot Racks

To facilitate accurate tracking of collected specimens, we recommend that you set up a blood collection tube rack with the set of draw tubes, pre-labeled with the provided participant ID labels. The tubes should be in the rack according to the order in which they are to be drawn, as specified above.

An aliquot rack, with pre-labeled cryovials, should be set up to correspond with each participant’s blood collection tube rack; and the cryovials should be in numerical order.  It may be helpful to have the red cryovials per participant in a separate rack since the red serum collection tubes are generally centrifuged at a different time from the other tubes.

### Preparation for Specimen Collection

Preparation for specimen collection is done in the following manner.  Early morning, prior to arrival of any participants:

1.  Make sure venipuncture supplies are stocked and the tubes and cryovials are labeled.

2.  Check that the sample processing station is properly equipped.  Every item on the checklist must be ready and in its proper position.

- Make sure the phlebotomy room is tidy and stocked with extra smelling salts, basin, washcloths, and that the draw tube mixer is functional.

- Label the tubes and cryovials with the participant ID (if not previously done).

- Approximately 10 minutes before scheduled blood specimen collection, fill styrofoam ice bath ¾ full with crushed ice.

5.6	Preparation of Participants – is similar to that of Exam 1 and 2.  A few points are worth re-stating here.

5.61	The participant’s experience must be as pleasant as possible. Give the participant enough time to feel comfortable, both before and after the blood collection. In many cases the most memorable part of the experience for the participant will be the contact with, and the attitude and competence of, the technician who draws the blood. Do not under any circumstances force or coerce the participant to have blood drawn.

5.62	Participants who are concerned about the volume of blood collected should be reassured that the total amount of blood drawn is about 4 tablespoons, although it may look like more. Additionally, reassure them that blood cells die and are made continuously and what is collected will be reproduced in one to two days.

5.63	Phlebotomy Form Questions.  There are five questions to ask the participant before the start of venipuncture.  The first three questions deal with the participant’s experience with venipuncture.  If they answer yes to any of these three questions, the phlebotomist can take extra care with the procedure.     Question 4 deals with diabetes status.  Check yes only if the participant is taking medication for diabetes.  Question 5 deals with fasting status.  The participant should be fasting (nothing to eat or drink except water) for 12 hours with a minimum acceptance of 8 hours.  If s/he have not fasted for at least 8 hours, the blood collection needs to be rescheduled.

5.7  	Venipuncture Procedure – is similar to that of Exam 1 and 2.

ALWAYS WEAR LATEX GLOVES AND LAB COAT

1.	Arrange draw tubes in order of draw on the table top or in the tube rack within easy reach. Assemble butterfly apparatus and Vacutainer holders, gauze, and alcohol prep prior to tourniquet application.

2.	Apply tourniquet.

- Examine participant's arms for the best site for venipuncture. Release tourniquet.

- Cleanse venipuncture site by wiping with alcohol prep pad in a circular motion from center to periphery.  Allow area to dry.

- Reapply tourniquet and start timer and document start time.  NOTE: If possible, it is best to release the tourniquet as soon as possible after flow has been established.  Tightened tourniquet should be on no longer than 2 minutes recommended or loosen tourniquet, then reapply if necessary. In our experience, however, especially with sick and/or elderly subjects, this may result in flow stopping, and the trauma of a second venipuncture. Therefore, this is a “judgment call” based upon the phlebotomist’s experience and skill.

6.	Grasp the participant's arm firmly, using your thumb to draw the skin taut. This anchors the vein.  The thumb should be 1 or 2 inches below the venipuncture site.

7.	With the needle bevel upward, enter the vein in a smooth continuous motion.

8.	Make sure the participant's arm is in a flat or downward position while maintaining the tube below the site when the needle is in the vein.  It may be helpful to have the participant make a fist with the opposite hand and place it under the elbow for support.

9.	Grasp the flange of the Vacutainer holder and gently push the tube forward until the butt end of the needle punctures the stopper, exposing the full lumen of the needle.

NOTE : Attention should be paid to minimizing turbulence whenever possible.    Small steps, such as slanting the needle in the Vacutainer to have the blood run down the side of the tube instead of shooting all the way to the bottom, may result in significant improvement.

10.	Note the blood flow into the first collection tube.  If blood is flowing freely, the butterfly needle can be taped to the participant's arm for the duration of the draw.  If the flow rate is very slow, the needle may not be positioned correctly.  Try moving the needle slightly without causing discomfort to the participant.

- Keep a constant, slight forward pressure (in the direction of the needle) on the end of the tube.  This prevents release of the shutoff valve and stopping of blood flow.  Do not vary pressure nor reintroduce pressure after completion of the draw.

- Fill each Vacutainer tube as completely as possible; i.e., until the vacuum is exhausted and blood flow ceases.  If a Vacutainer tube fills only partially, remove the tube and attach another of the same type without removing the needle from vein.  Tubes are not acceptable if < ½ full.

- When the blood flow ceases, remove the tube from the Vacutainer holder.  The shutoff valve re-covers the point, stopping blood flow until the next tube is inserted (if necessary). Place all tubes, except serum, on tube mixer for a minimum of 30 seconds.

- Release tourniquet, if still applied.   The ideal tourniquet time is two minutes.

- To remove the needle, lightly place clean gauze over venipuncture site.  Remove the needle quickly and immediately apply pressure to the site with a gauze pad. Have the participant hold the gauze pad firmly for one to two minutes to prevent a hematoma.   Discard needle into puncture-proof sharps container. Record on Phlebotomy form duration tourniquet was applied and end venipuncture time.

- The Citrate, EDTA, and SCAT-I, tubes are placed on wet ice.  The serum   tubes are held at room temperature.

- Clean up the venipuncture area (if necessary). Dispose of needle and tubing in the appropriate biohazard needle sharps containers. Complete the Phlebotomy Form.

18.	Bring the filled blood collection tubes to the processing area, keeping the EDTA, citrate and SCAT-I tubes on ice and the serum tubes at room temperature.

5.8	Guidelines for Difficulties – are the same as Exam 1 and 2.

Handling participants who are extremely apprehensive about having blood drawn.  Do not under any circumstances force the participant to have blood drawn.  It may help to explain to the participant that the blood drawing is designed to be as nearly painless as possible.  It is sometimes best to let the participant go on with another part of the visit.  It may also be helpful to have the participant relax in the blood drawing chair just so the phlebotomist can check the veins in the participant's arms, without actually drawing blood.  If the participant has "good veins" the phlebotomist can reassuringly say, "Oh, you have good veins; there should be no problem."

1.	Bandaging the Arm.  If the patient continues to bleed apply pressure to the site with a gauze pad.  Keep the arm elevated until the bleeding stops.  A gauze bandage can be tightly wrapped around the arm over the pad, and left on for at least 15 minutes.

2.	Procedures for Difficult Draw.  If a blood sample is not forthcoming, the following manipulations may be helpful.

a.  If there is a sucking sound, turn needle slightly or lift the holder in an effort to move the bevel edge away from the wall of the vein.

b.  If no blood appears, move needle slightly in hope of entering vein.  Do not probe.  If not successful, release tourniquet and remove needle.  A second attempt can be made on the other arm.

c.  Loosen the tourniquet.  It may have been applied too tightly, thereby stopping the blood flow.  Reapply the tourniquet loosely.  If the tourniquet is a velcro type, quickly release and press back together.  Be sure, however, that the tourniquet remains on for no longer than two minutes at a time.

d.  The phlebotomist should not attempt a venipuncture more than twice.

e.  Reassure the participant that the inability to obtain a clean venipuncture is not any sign of a medical problem on their part.

f.  If venipuncture is unsuccessful, this should be noted on the Phlebotomy Form.

3.	WHEN A PARTICIPANT FEELS FAINT OR LOOKS FAINT FOLLOWING THE BLOOD COLLECTION.

a.  Have the person remain in the chair, if necessary have him/her sit with head between knees.

b.  Provide the person with a basin if he/she feels nauseous.

c.  Have the person remain seated until he/she feels better.

d.  Place a cold washcloth on the back of the person's neck.

e.  If the person faints, use smelling salts to revive by crushing the ampoule and waving it under the person's nose for a few seconds.

f.  If the person continues to feel sick, contact a medical staff member who will advise you on further action.

##### Other Possible Problems:  Not all tubes are collected (blood flow ceases, difficult venipuncture, etc.).  Always fill collection tubes in the order specified.  Make notations of difficulties on the Phlebotomy form.  If the participant is willing, another attempt should be made to complete the draw collecting only those tubes that were not filled in the first venipuncture following the same tube order.

5.	Other Possible Problems:  Collection tube does not fill. First, try another tube of the same type.  Partially filled plasma tubes are not acceptable if less than ½ full.  If a tube is less than ½ filled, it should be discarded.  Partial tubes for serum are acceptable, but will result in a reduced number of aliquots.  If a tube is not completely filled (> ½), clearly note on the Processing Form as this can effect future assays.

6.	Processing Specimens

A.  Overview

Processing should be initiated as soon as possible (0 – 30 minutes) following venipuncture.  The red-topped serum tubes must stand at room temperature for at least 40 minutes before centrifugation. If centrifugation of the other tubes is not immediate, the citrate, EDTA, and SCAT-I tubes should remain on ice. Personal protective equipment (non-permeable lab coats, double-gloves with at least one latex pair, splatter shields are recommended) MUST BE worn for processing.

B.	Daily Preparation

The following items should be on hand before beginning processing:

•Lab coats and gloves, splash shields, other Personal Protective Equipment as needed.

•Refrigerated Centrifuge: 2,000 g-force minimum, 4 oC, Swinging bucket.

•10% bleach solution (or approved biohazard disinfectant)

•test tube holders (adapters) for centrifuges

•Harvard Trip Balance / Pan balance

•water bottles for balance

•freezer (-70C or colder)

•emergency eye wash station

•biohazard trash can, with large and small biohazard bags (biohazardous waste puncture proof containers)

•test tube racks / cryovial racks

•Fixed volume pipettes with tips (MLA) and adjustable pipettes (Rainin, Finn, etc) with tips.  Volumes needed to pipette:  0.5 ml, 1.0 ml (200 to 1000 ul)

•cryogenic vials* (0.5 ml and 1.5 ml)

•tube (Sarstedt 4.0 mL for urine)*

•cryovial labels (from the CC)

•Revco Boxes (#5954 and #5956) and dividers (9 x 9 and 7x7)

•Styrofoam/insulated shipping boxes

•Refrigerator- for storage of special blood tubes

- Can be a household fridge.

- Cannot be the same as food fridge.

•Labels/lab tape for reagents

•Sharpie pens

•ID labels for cryovials and freezer boxes

* = provided by CBAL

C.	Description of Aliquots

Aliquot Assignments:

| Collection Tube | Min Volume Needed after Centrifugation | Number of Aliquot Tubes | Color Code | Volume per Aliquot Tube |
| --- | --- | --- | --- | --- |
| #1: 10mL EDTA | 4.0mL | 7 cryovials | Purple | (Combine plasma from Tubes 1 &4 before aliquoting) 7 @ 1.0mL |
| #4: 10mL EDTA | 4.0mL |  |  |  |
| #2: 10mL Serum | 4.0mL | 7 cryovials | Red | (Combine sera from Tubes 2 & 5 before aliquoting) 1 @ 0.5mL 6 @ 1.0mL |
| #5: 10mL Serum | 4.0mL |  |  |  |
| #3: 4.5mL Citrate | 2.0mL | 4 cryovials | Blue | 4 @ 0.5mL |
| #6: 5mL SCAT | 2.0mL | 4 cryovials | Yellow | 4 @ 0.5mL |
| # 7: Urine | 3.0mL | 4.0mL tube | White | 1 @ 3.0mL |

### D.    ALIQUOTING: Citrate, EDTA, SCAT-1 and Serum tubes

1.	Aliquoting consists of removing the serum or plasma in small amounts (e.g.: 0.5ml) by pipette and placing it into the appropriate color-coded cryovials (provided).  Color-coding is predetermined and used to identify sample type such as citrated plasma vs. SCAT-I plasma, etc.  If color-coding is not available, the sample type must be clearly identified on the label.

2.	This process must be done while the tubes and cryovials are on ice (unless otherwise noted).

3.	When aliquoting serum and plasma, be careful not to disturb the top of the cell layer with the pipette tip, as this will result in platelet, white cell and red cell contamination.

4.	Use a new pipette tip for each draw tube.

5.	Once the sample is aliquoted cryovials should be immediately (< 10 minutes) frozen in an upright position at -70oC or promptly placed on dry ice for quick freezing.

After centrifugation, pool plasma or serum of like tubes from the same participant, (e.g.: EDTA plasma from tubes 1 & 4; Serum tubes 2 & 5).  A disposable transfer pipette may be useful in transferring the plasma or serum from the centrifuged blood collection tube into a 15 ml or similar ‘pooling tube’.  Make sure the pooling tubes are clearly labeled with ID#s.   From the pooled plasma or serum, now in the 15 mL tube, pipette the appropriate volume into each cryovial for that draw tube type (i.e.: EDTA plasma = 7 purple- capped cryovials).

If any tubes are accidentally mixed during pipetting so that plasma is contaminated with red cells, they may be recentrifuged.

Upon completion of the processing steps, aliquots must be frozen at -70C or below within 10 minutes, or place immediately on dry ice.   Make sure all cryovials and tubes are frozen in the upright position.

E.	Summary of Timing Issues

After blood drawing, time before centrifugation:

EDTA, SCAT-1, and Citrate:  store on ice; preferably < 15 minutes (maximum < 30 minutes) before centrifuging.

Serum:  store at room temp for at least 40 minutes, but < 90 minutes prior to centrifuging.

After aliquoting, ALL samples must be frozen within 10 minutes or placed immediately on dry ice.

Aliquot racks will be set up to correspond to each blood collection tube rack.  Rack setup is completed the previous day.  All tubes and vials are labeled with sample ID labels (if not previously done) and arranged in appropriate working order.

### Low sample volume

If there is insufficient sample of a tube type to make the full set of aliquots, if possible fill the cryovial that is marked with an * on the Processing page for that tube type first.

Any partially filled cryovial (less than the specified volume) should be marked with a dot on the cap and a “P” in the comment field on the Processing Form next to that cryovial number.

F.	Centrifugation – EDTA,  SCAT-1, & Citrate Tubes

If centrifugation is not immediate, tubes are stored upright on wet ice. Tubes are centrifuged at 4C at least at 2,000g x 15 minutes or 3,000g x 10 minutes for a total of 30,000 g-minutes. Maximum time elapsed before centrifugation is 30 minutes from time of collection.  Please note all start times on the Processing Forms.  Once centrifugation is complete, tubes are carefully placed on ice and are ready to aliquot.

1.  EDTA.  Place tube in the centrifuge.  After centrifugation, the EDTA plasma is pooled and aliquoted, by specified volume into cryovials #01 - 07.  These will be purple-capped cryovials.

### 2.  CITRATE.  After centrifugation, carefully pipette 0.5ml of this plasma into each of the cryovials # 8 – 11.  These will be blue-capped cryovials.

3.  SCAT-1.  After centrifugation, carefully pipette 0.5ml of this plasma into each of the cryovials #12 – 15.  These will be yellow-capped cryovials.

G.	Serum - Centrifugation

Allow serum tubes to clot for at least 40 minutes at room temperature (maximum time before centrifugation is 90 minutes).  These tubes are centrifuged at 4C at 2,000g x 15 minutes or 3,000g x 10 minutes for a total of 30,000 g-minutes.

After centrifugation is complete, pool the serum before aliquoting and place on ice.  Carefully pipette 0.5ml of pooled serum into cryovial # 16.  Pipette 1.0 mL of pooled serum into each of the cryovials #17-22.  These will be red-capped cryovials.  The remaining red cells in the serum draw tube (and the tube itself) can be discarded in the biohazardous waste.

# H.	Special Circumstances

1.  Blood specimens (EDTA, Citrate, SCAT-I) cannot be processed within 30 minutes of collection.

If centrifugation cannot be performed within 30 minutes of collection, try to process specimens as soon as possible after that time. Note time of collection and centrifugation on the P/P form. Maintain the EDTA, citrate, and SCAT-I tubes on wet ice until centrifugation.

2.  Serum and plasma cannot be frozen within 10 minutes of aliquoting.

Every effort should be made to freeze serum and plasma cryovials at -70C or below as soon as possible after aliquoting. If specimens cannot be placed immediately at -70C or below, they may be temporarily (< 2 hours) stored at -20C or placed on dry ice until transfer to -70C or below.  Dry ice is the preferred solution.

I.	Processing Completion.

The completed Phlebotomy/Processing Forms are kept in a temporary file.   Enclose copies of the Processing Forms with each shipment of samples to the Central Blood Analysis Laboratory. Upon receipt at CBAL, forms and samples are examined for monitoring/QC purposes.

Completed, frozen cryovials from three participants are packed into one freezer box.  All urine samples (#23) are packed into a larger size freezer box.

Be sure the Phlebotomy and Processing Forms are completely filled out.

## Wipe down all work areas with 10% Bleach solution (or approved biohazard disinfectant).

## Label and arrange cryovials in their proper racks for the next days blood processing

SHIPPING BLOOD SAMPLES

A. General

Blood samples are shipped only on Mondays or Tuesdays to the CBAL by an overnight carrier (Federal Express is preferred).  Samples will be shipped on a pre-arranged schedule.

B.  Packaging Samples

Sample Shipping Checklist:

Coolers

Rubber bands for freezer boxes

Ziplock plastic bags for freezer boxes

Absorbent material (i.e. paper towels, newspaper)

Packaging tape

Dry ice (~10lbs per mailing container)

Labels:  Fedex address labels,

Air eligible sticker (airplane)

Dry Ice Labels (class 9, UN1845)

Labeled freezer boxes with participant samples

Completed Processing Forms

Completed Shipping Forms (to be faxed)

C.  Procedure

For frozen shipment to the University of Vermont:

1.  Line cooler with absorbent material (i.e. paper towels).

2.  Place approximately 1/2 the dry ice (per mailer) on the bottom of the cooler.

3.  Place another layer of absorbent material (i.e.: paper towels) on top of the dry ice – so it will be between the dry ice and the freezer boxes containing the samples.

4.  Collect the freezer boxes containing samples to be shipped, and check the sample ID numbers against the Processing Forms for that shipment.  Each cryovials box contains the samples from 3 participants.  Urine tubes are stored in a taller box with a 7 x 7 grid.

5.  Place a rubber band around each cardboard freezer box containing samples before enclosing each box in a ziplock plastic bag.  Then carefully place these bagged boxes containing samples in the mailer.  The rubber band is important for aiding in the prevention of a cryovial spill; the ziplock bag & absorbent material are for compliance with commercial carrier specifications.

6.  Another layer of absorbent material is placed on top of the sample freezer boxes.

7.  The remaining dry ice is placed on top of this last layer of absorbent material.

8.  The top of the cooler is sealed on with tape.

9.  The Processing Forms for all samples included in that particular shipment, are placed on the top of the absorbent material before the top is securely taped closed.

10.  Affix shipping label(s).  Place the entire box in the refrigerator if pickup is not immediate. (Samples should not be on dry ice for > 24 hours).

Fill out the Shipping Log including the Fedex airbill #s and fax to the University of Vermont at (802) 656-8965.

This shipping protocol follows the procedures mandated by the International Air Transport Association’s Dangerous Goods Regulations-Packaging Instructions 650 and 904. Copies of these regulations are included with this MOP.

D.   Mailing Addresses:

University of Vermont

Department of Pathology

Colchester Research Facility, Room T205

208 South Park Drive, Suite 2

Colchester, VT 05446

Attn:  Elaine Cornell

(802) 656-8963

(802) 656-8965 Fax

QUALITY ASSURANCE

# A.	Overview of Field Center Monitoring

Quality assurance monitoring of the blood collection and processing protocols is important for the identification of any deviations from the standardized methods.  Differences in the manner of blood collection or processing could potentially create a statistically significant difference in assay results.  In order to prevent any sample associated problems, the CBAL has designed a system for monitoring the quality of blood collection and processing in each Field Center. The first component in the quality assurance program for Field Centers consists of the CBAL training course and certification process for each Field Center technician.  Other components of the program include maintenance of equipment check logs at each field center,  Field Center Supervisor checklist, review of Phlebotomy/Processing Forms by the CBAL, and analysis of problems associated with the phlebotomy.  Through the monitoring of these parameters, any systematic or random problems should be identified and appropriate corrective actions taken.

# B.	Field Center Technician Training & Certification

Standardization of venipuncture and blood processing procedures is of utmost importance for the quality of the blood samples and subsequent data analysis.  There will be a one time training session on blood collection and processing of the MESA samples.  The training session will present information relating to the collection of the blood sample (i.e.: infection control, safety precautions including OSHA regulations, handling equipment, venipuncture procedure and possible venipuncture problems), and proper processing procedures for the varied array of draw tubes, including centrifugation and temperature requirements, and aliquoting the multitude of corresponding color-coded cryovials.

Field Center Technician Requirements.

Prior clinical phlebotomy experience is mandatory for the Field Center technicians who will be performing blood collection for the MESA study.

Field Center Technicians should have read the MESA Manual of Operations before attending the CBAL training session.  Certification in MESA blood collection and processing is required before working with actual participants and blood samples.

## Field Center Technician Certification

Field Center technicians who attend the CBAL training session and successfully complete both the written and practical examinations will be certified in MESA blood collection & processing.   Once fully certified, this technician is qualified to certify other technicians at their site in the complete or partial process with final approval from the CBAL.

The steps for certification are:

1.  Read the Lab Manual of Operations.

2.  Observe the process performed by a certified technician.

3.  Successful completion of practical exam (using the Certification

4.  Form/Supervisor Checklist), which involves observation by a certified personnel of complete phlebotomy/processing procedure on a volunteer.

5.  Successful completion of written exam (prepared by CBAL).

Completed written exams will be corrected and kept on file at the CBAL.

# C.	Field Center Equipment Records

Each Field Center is responsible for maintaining daily and monthly records for equipment performance.  Daily temperature checks on refrigerators, freezers and refrigerated centrifuges should be performed.  Equipment temperature logs are filed on site for future reference and reported to the CBAL monthly.  These equipment records can identify problems with sample quality in the aliquoting and local storage steps.

- D.	Field Center Supervisor Checklist

The Field Center Supervisor checklist serves as a periodic monitoring measure.  The Field Center Supervisor will observe the MESA technicians at their site while they perform the phlebotomy and processing procedures, recording their observation on the checklist.  Completed Supervisor Checklists will be sent to the CBAL for monitoring purposes.   Checklists need to be completed once per month per technician.

E.		Maintaining Certification

A technician must perform phlebotomy and/ or processing on a minimum of one participant, every two weeks in order to maintain certification.

- Field Center Acknowledgement Forms

The Vermont Central Blood Lab will analyze the condition of each shipment received and will complete an Acknowledgement Form and fax it back to the Field Centers.   The Acknowledgement Form is used as a tool to track possible problems and variations from protocol on a weekly basis.

Alert Values

The University of Minnesota will be analyzing these samples for Lipid Panel, Glucose, and Creatinine.   The University of Vermont will run microalbumin on the urine samples.

Alert values are as follows:

Total Cholesterol		>360 mg/dL

Triglycerides			>1000 mg/dL

HDL cholesterol		< 20 mg/dL

LDL cholesterol		> 260 mg/dL

Glucose 			< 50 or > 400 mg/dL

Creatinine			> 2.0 mg/dL

Blind Duplicates

Blind duplicate samples will be collected on 5% of the participants on four different tube types.  This results in 20% of all participants having a blind duplicate sample collected for QC.  The criteria for collecting a blind duplicate sample will be based on a check digit in the participant’s ID number.

The following tube types will be drawn (only one tube per participant depending on the check digit):

5 ml EDTA (B-D# 366452)

5 ml Serum (B-D# 366534)

4.5 ml Citrate (B-D# 366415)

5 ml SCAT-I*

*Provided by CBAL

The blind duplicate tube is collected after the regular tubes are filled.  It would be the seventh tube filled.

The tubes are handled in the same way as the regular collection tubes.   EDTA, Citrate, SCAT-I are placed on the mixer for approximately 30 seconds, then placed in ice, and centrifuged within 15 to 30 minutes.  Serum remains at room temperature for a minimum of 40 minutes to clot before centrifuging.

Aliquoting Scheme:

# Tube Type				Cryovial color		# x  sample volume

EDTA					purple				2 x 1.0 ml

Serum					red				4 x 0.5 ml

Citrate					blue				4 x 0.5 ml

SCAT-I				yellow				4 x 0.5 ml

Cryovials must be labeled with a QC ID#.   This ID# is matched to the Participant ID#.

After aliquoting, cryovials are frozen immediately at –70o C in an upright position.     Blind Duplicate cryovials are placed in their own freezer box with a 9 x 9 grid.   More than one participant’s samples are included in one box.  These samples are shipped a week or so after the original samples are sent out to the CBAL, so that the laboratory cannot match them with the original participant.   Include the completed Blind Duplicate Shipping Log when these frozen samples are shipped.

# Urine Blind Duplicates

Blind duplicates will also be processed for urine.  Whenever an EDTA blind duplicate is being collected on a participant, a urine blind duplicate should also be processed.  On each participant selected for EDTA QC, the QC ID# will be the same for both the EDTA and the urine blind duplicate cryovials.

After aliquoting the participant’s urine into tube #23, aliquot 3 ml of the remaining urine into a 4 ml tube.  Carefully label this tube with the QC ID# and freeze immediately.  Urine blind duplicates should be stored separately from the participant’s original sample set, in a freezer box with a 7 x 7 grid.

It is important to complete the Blind Duplicate Shipping Log and include a copy in the shipping box with the frozen samples.  Ship frozen urine blind duplicates at least one week after the original participant’s samples are shipped to CBAL.

Sample		Cryovial			# x sample volume

## Urine		tube 4.0 mL			1 x 3.0 ml

FIELD CENTER FORMS

MESA Phlebotomy/Processing Form

MESA Shipping Log (same as Baseline)

MESA Blind Duplicate Shipping Log (same as Baseline)

MESA Field Center Supervisor Checklist

MESA Field Center Technician Certification Examinations

MESA Equipment Temperature Logs (same as Baseline)

DIAGRAMS/INSTRUCTIONALS:

Aliquoting Scheme Flow Chart

Freezer Box Diagram

IATA Packing Instructions 650 and 904

3.8   Carotid Ultrasound

3.8.1 General Guidelines

## Standardization and Probe Frequency

MESA ultrasound scans will all be performed with GE Logiq 700 ultrasound machines using ML probes. The probe operates at two different frequencies: 9 and 13 MHz. Images of the brachial artery and common carotid are captured at 9 MHz, and the internal carotid images are captured at 13 MHz. Pulse Wave Doppler measures are not affected by the probe frequency setting. When making Pulse Wave Doppler measurements, the probe frequency can remain set at the frequency that makes the execution of the protocol most efficient.

## Data Transmission

Ultrasound scans are recorded to Super VHS videotape and sent to the Ultrasound Reading Center for analysis. Field Centers send tapes to the Reading Center every Thursday. Two separate tapes are used, one for both Brachial Artery Endothelial Function and Carotid Distensibility scans and the other for Carotid IMT scans. The separate tapes are used, because Endothelial Function and Carotid Distensibility analysis involves the capture of video loops, whereas Carotid IMT analysis involves static images. Since the Brachial Artery Endothelial Function and the carotid scanning are done at separate clinical sites at Columbia University (MESA Field Center 4) this field center uses three separated videotapes.

Forms

There are two different forms that the MESA sonographers will be using. One is the IMT Videotape Log Sheet and the other is the Ultrasound IMT Case Report Form (titled Carotid Ultrasound).

A videotape log sheet, which can consist of more than one page, is the reader’s guide to the scans that are recorded on the videotape. It is sent to the Ultrasound Reading Center with the videotape. Photocopy the log sheets before sending them.

MESA ultrasound case report forms are completed for each participant. These forms are scanned into the MESA database at the field center. Do not send these forms to the Ultrasound Reading Center.

Ultrasound Videotape Log Sheets and Labels

A videotape log sheet is to a MESA ultrasound videotape what a table of contents is to a book. Just as it is impossible to find out what is in a book without a table of contents, it is also impossible to determine what is on the tape and to analyze it without a log sheet. The process is simply much more efficient with the log sheet.

Matching numbered labels for both the log sheets and videotapes have been provided to facilitate the long-term organization of MESA ultrasound data. A new log sheet is started for each new videotape. Label each new videotape and log sheet with matching labels before scanning. If more than 10 participants will be recorded on videotape, attach a second page and continue.

Though most of the entries to the Videotape Log Sheet are self-explanatory, take the time to review the details below. It is important to check subject IDs and to write legibly. Transposition errors can be difficult to catch after the participant has left the scanning area.

VCR Start Time

The one item on the log sheet that can cause confusion is the “VCR Start Time.” The “VCR Start Time” recorded on the log sheet is analogous to the page numbers in a table of contents. Without the “VCR Start Time,” finding a study on a videotape can be very difficult and time consuming.

### General Guidelines for “VCR Start Time”

- Start each new tape with the counter set at 0:00

- At the end of each study mark down the VCR counter time of the study just completed as the "VCR Start Time" for the next study on the log sheet.

- If the next scan is immediately consecutive and on the same videotape, great. Start the next scan and when it is complete, mark down the VCR counter time as the "VCR Start Time" for the next study.

- If the scans are not consecutive, there are two options:

Option A (more efficient)

1.	Don't rewind the videotape, or, if you have reviewed, have it queued up (ready) at the end of the last recorded scan.

2.	Re-set the VCR counter to 0:00.

3.	Videotape the new study.

4.	When the scan is completed, jot down the time displayed on the VCR counter. Add it to the “VCR Start Time” of the scan just completed.

5.	Write the result in the “VCR Start Time” space of the next scan.

Option B

1.	Rewind the tape

2.	Reset the VCR Counter to 0:00.

3.	Fast-forward to the end of the last study.

4.	The VCR counter should now display at time within seconds of the time written in as the "VCR Start Time" on the log sheet.

5.	Videotape the new study.

6.	When the scan is completed, record the VCR counter time at the end of the study just completed as the "VCR Start Time" of the next study.

Videotape Log Sheet Details

| Item | Description |
| --- | --- |
| Scan Date | ### Date scan is performed |
| Subject ID | Participant’s MESA ID |
| Plaque Location #### Carotid IMT log sheet only | If there is (are) plaque(s), indicate the location(s) by circling the appropriate acronym. RB = right bulb RICA = right internal carotid LB = left bulb LICA = left internal carotid |
| VCR Start Time | VCR Counter time, not the time of day. Start new videotapes at 0:00. Re-set the VCR counter if necessary. |
| Sonog ID | Sonographer’s MESA Technician ID |
| ## Comments | This space has been provided for general comments regarding each scan. Note any significant findings, including plaque, wall thickening, or difficulties encountered in trying to obtain good images. This section of the log sheet is invaluable to the ultrasound Reader, who must rely on the sonographer to communicate what is seen. |

Sample Carotid IMT Videotape Log Sheet

Only Carotid IMT scans are recorded on the Carotid IMT videotape.

|  | Scan Date | Subject ID | Plaque  | VCR Start Time | Sonog. ID | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| 1. | 6/13/2000 | 3551212 | RB RICA LB LICA | 0:00 | 317 | Nice clean and easy scan |
| 2. | 6/13/2000 | 3899223 | RB RICA LB LICA | 3:10 | 319 | Plaque in right ICA |
| 3. | 6/14/2000 | 3221555 | RB RICA LB LICA | 5:25 | 317 | Use second LCC image, thick neck difficult to scan |

3.8.2 Right Common Carotid Dynamic Acquisition

## MESA Case Report Forms

All MESA case report forms require the entry of the MESA participant ID, acrostic and date at the top right, and the sonographer, reviewer and data entry IDs along the bottom.

Protocol: Carotid Progression and IMT

Form Name: Carotid Ultrasound

| Item | Response Options | Explanation |
| --- | --- | --- |
| ID # | primary MESA Subject Identifier |  |
| Acrostic | secondary MESA Subject Identifier |  |
| Date | scan date – today’s date |  |
| - Results of Carotid IMT Progression Examination (20 seconds of right common carotid dynamic acquisition): | - Done - Incomplete - Not Done | Indicate whether or not the scan was successfully conducted and recorded. Skip to #3 if it was “Done.” |
| - Reason Carotid IMT Progression Examination incomplete or not done: | - Equipment Malfunction - Time/Staff/Room Constraints - Examinee refused/uncooperative - Examinee physically unable - Other (specify) | If scan is “Incomplete” or “Not Done,” indicate why. |
| - Tape # | 3 digit number | The number on the videotape and log sheet. No letters. |
| - VCR Start Time | 1 digit for hour 2 digits for min 2 digits for sec | VCR Counter time, not the time of day. |
| - Were Right CCA Doppler blood flow signals detectable? | - Yes - No | Indicate whether or not PW Dopplers could be heard. |
| - Right CCA Pulse Wave Doppler measurement (cm/s) | 3 digits cm/sec | Enter PW Doppler measurement from image screen in cm/sec. |
| - Exam 1 image evaluation | - None available - Device malfunction – not available - Landmarks identifiable? If Yes for “Landmarks,” select one: - All - In part - None | Skip to #10 if “None available” or “Device malfunction—not available” is marked. If “Landmarks” is marked, then also fill in “All,” “In part,” or “None.” |
| - Quality of Exam 1 image | - Excellent - Very Good - Satisfactory - Poor |  |
| - Quality of current examination | - Excellent - Very Good - Satisfactory - Poor |  |
| - Post-imaging Blood Pressure and Pulse | 3 digits for systolic 3 digits for diastolic 3 digits for pulse | Left arm blood pressure (use right arm if there is a gradient ≥ 15 mmHg; check box if right arm used) |
| - Results of Carotid IMT scan | - Done - Incomplete - Not Done | Indicate whether or not the scan was successfully conducted and recorded. Skip to #13 if it was “Done.” If “Not Done,” then complete Question 12 and skip to end of form. |
| - Reason Carotid IMT scan incomplete or not done | - Equipment malfunction - Time/Staff/Room constraints - Refused/Uncooperative - Examinee physically unable - Other (specify) |  |
| - Were ICA Doppler blood flow signals detectable? (Right and Left) | - Yes - No | Indicate separately for the right and the left sides whether or not PW Dopplers could be heard. |
| - Pulse Wave Doppler measurements (cm/s) (Right and Left ICA/Bulb) | 3 digits cm/sec | Enter PW Doppler measurement from image screen in cm/sec. |
| - Quality of scan (Carotid IMT scan) | - Good - Fair - Poor | Based on your assessment, is the quality of this Carotid IMT scan good, fair, or poor? |

Carotid Imaging Protocol

The protocol includes review of previous common carotid artery images taken during Exam 1 on the participant.  This Exam 1 image is distributed on a CD-ROM with the subject ID as an identifier. A reading program included on the CD permits image display on any standard Window PC.  This image is used to confirm orientation and location of the previous study and serves as an aid for the sonographer in localizing the bulb and other landmarks.

The common carotid artery (CCA) dilates to form the carotid bulb proximal to the bifurcation into the internal and external carotid arteries. The origin of the bulb can be recognized in most, though not all, subjects. The common carotid artery bulb is defined as the site where the artery begins to dilate slightly and the vessel walls curve out, i.e.: they are no longer parallel to each other. The bulb is elliptically shaped and geometrically complex in the longitudinal view. The tip of the flow divider defines the bulb’s upper limit. The tip of the flow divider also marks the origin of the internal and external carotid arteries.

Evaluation of the common carotid artery IMT is made on a segment at least 1 cm below the carotid bulb. The common carotid artery is, at this level, straight. Blood pressure measurements will be made to verify that there are no significant interval changes in blood pressure since the previous carotid IMT examination.  Variations in blood pressure may affect IMT estimates.  For example, an increase in blood pressure will increase the diameter of the common carotid artery.  The short-term effect of a diameter increase of the common carotid artery is to slightly decrease the IMT at the level where the IMT measurement is being made.

## Videotape Identification – Labels and Log sheets

The Ultrasound Reading Center will provide videotape labels and log sheets to the field centers. Matching numbered labels for the videotapes enable the long-term organization of MESA data. A new log sheet is started for each new videotape. Before beginning the first scan, attach the matching labels to the videotape and log sheet. Labels are put on the spine and the face of the videotape and the matching label is put in the specified area of the log sheet.

MESA Exam 3 Carotid Ultrasound procedure form

A procedure form will be completed for every participant by the sonographer. The forms are scanned into the MESA database. In addition, to facilitate the analysis process, photocopies of the form are to be made and sent to the Ultrasound Reading Center together with the videotape and videotape log sheet.

## Participant Position and Blood Pressure Measurement

The subject is supine during the carotid artery examination and is made comfortable in a position that allows head rotation to the left side. The sonographer stands or is seated at the end of the examination table near the participant's head. The top of the head is about three inches from the end of the examination table and the head is rotated 45 degrees to the left. Participants may wish to turn their heads to look at the screen. This should be discouraged. The sonographer may gently place his/her hands on the subject’s head as an aid to rotate the head in the correct position. The sonographer should emphasize the importance of not moving. A blood pressure cuff is connected to an automated sphyngomanometer (Dinamap) placed over the left upper arm. Blood pressures are recorded from the left arm. The left arm is held slightly abducted from the body. Record the blood pressures and pulse to the MESA Exam 3 Ultrasound IMT procedure form.

Image review

Exam 1 common carotid images will be distributed to the field centers. They are in a standard file format that can be recognized by most digital imaging software including “Imaging” a program installed as part of the Windows operating system. The image file names include the MESA participant’s subject ID numbers. For example, the digital image for the Subject ID# 5021399 is 5021399.DIB.

The images are displayed by double clicking on the file name to open it or by starting the program “Imaging” and selecting File  Open and navigating to the appropriate file. Sonographers should verify that the MESA Subject ID imprinted upon the displayed Exam 1 CCA image is that of the participant being scanned. If the image file cannot be found or the imprinted Subject ID is does not match, please notify the Ultrasound Reading Center immediately.

Machine Set-up

All images are taken only from the right common carotid artery with the Common Carotid Presets entered in the GE Logiq 700 device.

Common Carotid Presets

Frequency: 13 MHz

66 Dyn Range

Res

Edge 3

Map G

Ave. 2

Number focal zones: 2  (position of focal zones: mid artery)

## Videotape Demographic Information and Label Images with Study Identifiers

Participant information is entered on the ultrasound machine’s demographic information screen to identify the scan as a MESA Exam 3 scan for a particular subject. This page is videotaped before scanning begins to ensure the identification of the scan data with the proper participant. Videotape the demographic information page for five seconds before beginning the imaging protocols. The following information should be included on the demographic screen and on the image throughout the scan:

- The study name and the exam: MESA Exam 3

- Subject ID

- Subject Acrostic

- Date of ultrasound scan

- Sonographer ID

- Protocol: IMT Progression

Enter subject information to subject information screen

The participant ID number is entered

Press "New patient"

Hit "return" until you reach the new patient question (Y/N); type "Y" and then return

Keep hitting return until the Last Name field: Type "right cca"

Keep hitting return until the ID# field: enter the ID number

Hit return for gender (M or F)

Hit return and enter your Sonographer ID

###### Videotape screen for 5 seconds

Hit the "Exit" key to exit to begin scanning

Initial Scan (Videotaped for 10 seconds at a 4 cm field-of-view)

The purpose of the initial scan is to orient the sonographer to the subject’s carotid anatomy:

- To identify the relative location of the jugular vein to the common carotid artery,

- To locate the bifurcation and;

- To locate the common carotid artery bulb.

- To confirm proper orientation of the transducer

Videotape ON

This initial scan is done in the plane transverse to the artery and vein.  The transducer is slowly swept from the low common carotid artery (just above the clavicle) to just above the origins of the internal and external carotid arteries: the transducer is kept over the internal jugular vein so that the vein is stacked on top of the artery. Once the location of the bifurcation is identified, the transducer is slowly rotated to a plane parallel to the axis of the artery and vein. The jugular vein then serves as an acoustic window for imaging of the common carotid artery. Apply only a minimal amount of pressure to the transducer in order to prevent collapse of the vein and loss of the acoustic window.
Videotape OFF

Summary of imaging sequences:

- Video stream 1: A pulse-wave Doppler tracing is recorded over at least 5 cardiac cycles, approximately 5 seconds of real-time imaging.

- Video stream 2: A recording is made with the optimal "zoomed" image of the common carotid artery. Both near and far wall interfaces defining the intima-media thickness should be clearly imaged over at least a 1 cm length. This recording is made for 20 seconds (approximately 20 cardiac cycles)

Imaging sequences:

Video Stream 1 - Pulse-wave Doppler

A 4 cm field-of-view is used as the default image size.

The critical information is the peak-velocity in the common carotid artery at peak systole at the point of maximum flow acceleration. This confirms location and patency of the carotid artery.

Procedure:

- Hit the "cursor" key.

- Move the cursor to the middle of the artery with the trackball

- Press the toggle switches ("angle steer" and "angle correct") until there is an angle of 60 degrees or less between artery axis and the ultrasound beam.

- Press the PW button. Adjust the velocity scale so that the Doppler waveform fits the display (Y-axis).

- Turn the videotape on for 5 -10 seconds until a satisfactory Doppler tracing is obtained

- Freeze the image and let the videotape record the frozen image for 5 seconds

N.B. The velocity scale is set so the Doppler waveform fills the display.

- Record peak Pulse Wave Doppler measurement to MESA Exam 3 Ultrasound IMT procedure form.

Video Stream 2 - Right Common Carotid real-time gray scale imaging

Verification of image location:

If it is available, review the appearance of the Exam 1 CCA image on laptop computer screen. Note the location and orientation of the bulb, artery wall and if visible the tip of the flow divider. Document the visibility of the landmarks and assess the quality of the Exam 1 image on the procedure form.

## Image acquisition

Imaging is done in the lateral projection with the jugular vein lying immediately above the common carotid artery (or at 45 degrees if the internal jugular vein is not present). The image is centered on a 10 mm segment of the right common carotid artery at least 10-mm below (caudad to) the right common carotid artery bulb. The carotid bulb may be displayed on the left side of the monitor (when facing the screen). The bulb occupies 1/4 up to 1/3 of the image along the X-axis.  If the bulb cannot be identified, but the tip of the flow divider can, this may substitute as the internal landmark on this view.

After locating the tip of the flow divider on the transverse image, the transducer is slowly moved down the neck by 2 to 3 cm.  Rotate the transducer into the lateral plane keeping the jugular vein in the same imaging plane. The probe is then centered on the upper 2 cm of the common carotid.

The sonographer magnifies the longitudinal image (jugular vein above common carotid artery).

- Press the "ZOOM" button

- Use the trackball to center the distal common carotid artery (edge of bulb occupying up to 1/3 of the image to the left hand side of the screen)

- Press the "SET" button

- A videotape is taken for a minimum of 20 seconds so that the images in the video sequence are from at least 20 cardiac cycles. The common carotid artery wall interfaces (near and far wall IMT interfaces) are clearly depicted during this recording.

## Procedure Form Details

A MESA Exam 3 Carotid Ultrasound procedure form will be completed by the sonographer for every MESA subject scanned. The sonographer is responsible for the numbered items and to verify the Subject ID, Acrostic and the Date. New Procedure: the procedure forms are to be photocopied and the copies sent to the Ultrasound Reading Center together with the videotape and the log sheet. The originals are to be returned to the clinic coordinator who will oversee their being scanned into the MESA database system.

##### Summary of MESA Exam 3 Procedures

| ##### Preparation |  |
| --- | --- |
| ##### Videotape | ##### Labeled and queued up |
| Log sheet | Labeled and completed |
| Exam 1 Image Review | Locate and display Exam 1 image on laptop |
| Ultrasound IMT procedure form | Verify Subject identifiers, date, Sonographer ID |
| Pre-sets – Common Carotid | Verified |
| ## Protocol |  |
| Blood pressure | Measure with Dinamap and record on form |
| Demographics | Enter Subject and scan information ## Videotape for 5 seconds |
| ##### Transverse sweep | ##### Videotape 10 second transverse sweep from the low common to above origins of ICA and ECA with the jugular vein stacked on top of artery |
| Video stream 1 (PW Doppler - CCA) | 4 cm field of view Videotape 5-10 seconds of live Doppler tracing Videotape 5 seconds of measured peak velocity Record peak PW velocity on form |
| Video stream 2 (Longitudinal gray scale of CCA) | ZOOMED field of view Match the Exam 1 image Videotape 20 seconds of right CCA: clear near and far walls |

Data Transmission

Videotapes are sent weekly to Dr. Polak through a package delivery service with tracking capabilities at the following address. The package includes the appropriate videotape and log sheet and copies of the procedure forms.

Joseph F. Polak, M.D., M.P.H.

MESA Ultrasound Reading Center

New England Medical Center, Box 380

750 Washington Street

Boston, MA 02111

##### Carotid Acquisition/Analysis quality assurance

All new MESA sonographers will travel to Boston for a one-day training session.  This training session is similar in content to the MESA training session and consists of a review of basic carotid physiology and instrumentation.  Following didactic lessons, the sonographers will practice imaging according to the protocol on volunteers, using the GE Logiq-700 ultrasound device available at the Field centers. Dr. Polak will be in charge of this session. Following this training session, the sonographers will conduct 10 examinations in order to be certified.  The certification scans will be reviewed for protocol adherence and quality and a written critique provided. Once certified, they will start imaging the MESA cohort.

Sonographers who have been certified by the Ultrasound Reading Center during Exam 1 or Exam 2 are eligible to scan for Exam 3 after participating in a telephone conference call with Dr. Polak and the URC staff. The sonographers will have hardcopies of the manual, procedure form and videotape log sheet, which will be distributed via the Internet, with them for reference during the call. The protocol will be discussed, administrative details will be reviewed and questions will be answered. The first tapes will be carefully reviewed for protocol adherence and image quality and a written critique provided to each sonographer.

All readers will need to be certified.  This will be done after training (one month of performing IMT measurements).  Dr. Polak will review at least 30 common carotid IMT measurements in order to certify the readers.

Quality assurance will be conducted in order to determine inter-sonographer (if applicable), intra-sonographer variability, intra-reader variability and inter-reader variability.

Since new sonographers may be working at some of the Field centers, we need to make allowances for differences in sonographer performance.  This will done by reader evaluation of the relative location/presence of the carotid bulb on the current study image as compared to the Exam 1 image.  The reader will score concordance of location for the current study image, as compared to the Exam 1 image, on a 3 level Likert scale (exact concordance, partial concordance, no concordance).  The reader will also score concordance of image gain for the current study image, as compared to the Exam 1 image, on a 3 level Likert scale (exact concordance, partial concordance, no concordance). The reader will also judge the quality of the Exam 1 study image and the current study image according to a five level Likert scale: (perfect alignment, almost perfect, interpretable, poorly interpretable, unintepretable).

We will perform intra-sonographer variability evaluations at the beginning of the examination, mid-way into the examination and at the end of the examination.  Sonographer will repeat 30 ultrasound image acquisitions at each time point.  At those centers where there are more than one sonographer, we will ask the sonographers to perform, in the first month of the study, separate image acquisitions on all these same participants (estimated 30 studies).  This will serve as a baseline for inter-sonographer variability.  This will also be repeated at 9 months into the examination and during the last month of the examination.  These replicate scans will be measured by two different reader in order to determine (using a nested design) the relative contributions of inter-sonographer, intra-sonographer and, if applicable, inter-reader variability.

The consistency in subjectively evaluating the quality of the images and concordance between image pairs will be evaluated by weighed kappas (expected level of 0.6 or better).  Errors in paired measurements will be quantitated using paired absolute differences in IMT values (inter-reader, intra-reader, intra-sonographer and inter-sonographer, if applicable) in addition to ANOVA.

Since baseline IMT may turn out to be an important co-variate (see data analysis section below), appropriate corrections will be needed to account for bias between different readers.  Variability of human drawn interfaces and edge detector interfaces will be determined for the whole cohort.  Both types of IMT measurements will be available for estimating the extent of atherosclerosis progression.

3.8.3  Carotid IMT

I.	Overview of the Carotid Ultrasound Scanning Protocol

1.	Introduction to the Carotid Ultrasound Scanning Protocol

2.	Participant Position

3.	Anatomical Sites

4.	Initial Scan

5.	Standard Carotid Ultrasound Images

6.	Criteria for Satisfactory Images

7.	Imaging Priorities

8.	Sonographer Response to a Significant Stenosis – ALERT

9.	A Good Carotid Ultrasound Image

II.	Overview of Carotid Ultrasound Analysis Protocol

1.	Preview and Digitization of a Carotid Ultrasound Study

2.	Calibration

3.	The Interfaces

4.	Priorities

5.	Quality Scores

6.	Subjective Assessment

6.1	Surface

6.2	Morphology

6.3	Percent Stenosis

6.4	Location

6.5	Density

III.	Data Transmission

1.	Videotape

2.	Labels and Log Sheets

3.	Shipping Videotapes and Log Sheets to the Ultrasound Reading Center

4.	Data Transmission to the Coordinating Center

IV.	Sonographer Quality Control

1.	Clinical Center Sonographers

2.	Ultrasound Reading Center – Sonographer Communication

3.	Ultrasound Reading Center Report on Sonographer Performance

4.	Performance Standards during Examination Period

5.	Replacement or Retraining of a Sonographer

6.	With-in Sonographer Variability Assessment

7.	Achieving and Maintaining Quality Ultrasound Data

I.	Overview of the Carotid Artery IMT Ultrasound Scanning Protocol

1.	Introduction to the Carotid IMT Ultrasound Scanning Protocol

The extracranial carotid arteries are the largest arteries in the neck. The right common carotid artery (CCA) originates from the innominate artery on the right, and the left CCA originates directly from the aortic arch. Each common carotid artery ascends in the neck, lateral to and posterior to the trachea. At the approximate level of thyroid cartilage, slightly below the angle of the mandible, the common carotid artery bifurcates into the external and internal carotid arteries. Proximal to the bifurcation, the common carotid artery dilates to form the carotid bulb. The origin of the bulb can be recognized in most, though not all, subjects. It is defined to be where the common carotid artery begins to dilate slightly and where the vessel walls curve out and are no longer parallel to the skin surface. The bulb is elliptically shaped and geometrically complex in the longitudinal view. Its upper limit is defined by the tip of the flow divider. The tip of the flow divider also marks the origin of the internal and external carotid arteries. The external carotid artery lies anterior and slightly medially to the internal carotid artery in 90% of individuals. In the remaining 10%, the orientation is reversed. The external carotid artery is usually smaller than the internal and it has branches that supply the neck and face. The internal carotid artery has no branches in the neck and ascends into the calvarium to supply the brain.

2. 	Participant Position

The subject is supine during the carotid artery examination and is made comfortable in a position that allows head rotation to either side. The sonographer stands or is seated at the end of the exam table near the participant's head. The top of the head is about three inches from the end of the exam table and the head is rotated 35o away from the side examined. Participants may wish to turn their heads to look at the screen. This should be discouraged. This is best achieved by the sonographer using his hands at the start of the study to rotate the head while stressing to the participant the importance of not moving.

3. 	Anatomical Sites of Interest

3.1	The extracranial carotid arteries are divided into four anatomically defined segments:

	Distal common carotid artery

	Carotid bulb

	Internal carotid artery

	External carotid artery

3.2	The lateral extent of each segment is defined relative to the tip of the flow divider, which is typically the most clearly defined anatomical reference in the carotid system. The three segments of interest are the distal common carotid artery, the carotid bulb, and the internal carotid artery. No external carotid images will be recorded.

3.3	Anatomical Definitions

	Distal common carotid is the segment of the common carotid artery immediately proximal to the origin of the carotid bulb, where the near and far walls of the artery are parallel to one another. The end of the distal common carotid artery is marked by the dilatation of the vessel walls, which is the carotid bulb.

	Carotid bulb: the inferior extent of the bulb is the beginning of dilatation, or 8 mm below the tip of the flow divider. The superior extent of the bulb is defined by the very tip of the flow divider.

	Internal carotid artery: the caudal, or inferior, extent is defined by the tip of the flow divider. The vessel then ascends in the neck and enters the base of the skull. For the purposes of this protocol, the ultrasound study will be limited to the initial 10 mm of the internal carotid artery.

4. 	Initial Scan

The purpose of the initial scan is to orient the sonographer to the subject’s carotid anatomy. The sonographer should locate the bifurcation and distinguish which vessel is the internal and which is the external carotid artery. The site of maximal wall thickening in the near or far wall, in the bulb or internal carotid artery should also be identified during the initial scan. Color and pulse-wave Doppler can be used as identification aids. The initial scan is not recorded.

5.	Standard Carotid Ultrasound Images

5.1	Summary overview: one videotaped transverse (short-axis) scanning sequence and five videotaped longitudinal images are taken from both the right and the left carotid arteries for each subject. The transverse (short-axis) scanning sequence is a sweep of the carotid from the base of the common up through the bulb into the internal and back down to the base of the common. The purpose of the transverse scanning sequence is to help the reader distinguish between real echoes that represent plaque and spurious echoes that represent artifact. Occasionally when viewing the transverse images the reader is uncertain whether echoes represent real plaque or not. The transverse view serves as the tiebreaker. The first of the five longitudinal images is a pulse-wave Doppler measurement of the peak systolic velocity in the internal carotid. Next is a standard view of the common carotid. The other three longitudinal images are of the internal carotid artery or bulb. The frozen images of the common carotid and the internal or the bulb are videotaped as both frozen images and as part of a cine loop.

5.2	The magnification function is used when obtaining the lateral the common carotid and the internal carotid images. It is turned “ON” immediately after the pulse-wave Doppler measurement and before imaging the common carotid. The magnification must be kept constant throughout the remainder of the scanning on each side. This is because the reader calibrates on the first images and assumes that the magnification remains constant. An unexpected magnification change can result in erroneous measurements.

5.3	In summary, the sonographer first videotapes a transverse (short-axis) scanning sweep of the common through the bulb. He/she then finds the site of peak systolic velocity as measured by pulse-wave Doppler. The common carotid is then imaged. And finally the site of maximal thickening in the internal artery or in the bulb is then imaged from three different angles, anterior, lateral and posterior, to obtain as much circumferential information at this site as possible.

5.3	Collect and videotape the carotid ultrasound images in this order:

| Image Description | To Be Videotaped |
| --- | --- |
| Right Side |  |
| Standard zoom-depth = 4 cm |  |
| - Transverse (Short-axis) sweep | - 15 second scanning sweep |
| - Pulse Wave Doppler ICA | - 5 seconds of frozen, measured image |
| - Common Carotid | - 5 seconds of frozen image followed by the cine loop |
| - Anterior ICA or Bulb | - 5 seconds of frozen image followed by the cine loop |
| - Lateral ICA or Bulb | - 5 seconds of frozen image followed by the cine loop |
| - Posterior ICA or Bulb | 5 seconds of frozen image followed by the cine loop |
| Magnification “OFF” |  |
| Left Side |  |
| Standard zoom-depth = 4 cm |  |
| - Transverse (Short-axis) sweep | - 15 second real time scanning sweep |
| - Pulse Wave Doppler ICA | - 5 seconds of frozen, measured image |
| Magnification “ON” |  |
| - Common Carotid Artery | - 5 seconds of frozen image followed by the cine loop |
| - Anterior ICA or Bulb | - 5 seconds of frozen image followed by the cine loop |
| - Lateral ICA or Bulb | - 5 seconds of frozen image followed by the cine loop |
| - Posterior ICA or Bulb | 5 seconds of frozen image followed by the cine loop |

5.3.1	Image 1: Transverse (Short-Axis) Sweep

The transverse sweep begins at the base of the common carotid and smoothly travels up through the bulb, into the internal and back down to the common. Videotape in the full sweep, up from the common, into the internal and back, for 15 seconds. The transverse scan is done with the ultrasound machine in standard zoom mode at a depth of 4 cm.

5.3.2	Image 2: Pulse-wave Doppler

The critical information is the peak velocity in the internal carotid artery at peak systole at the point of maximum flow acceleration. To locate it a 2-mm Doppler sample gate should first be place in the center of the distal common carotid artery. The sample gate is then moved from the common through the bulb to the proximal internal. If there is no site of disturbed or turbulent flow, the Doppler measurement should be taken from the first centimeter of the internal carotid artery. The audible signal can be used to facilitate placement of the Doppler sample gate. Angle correction must not exceed 60o. The frozen image of the measurement is videotaped for approximately five seconds. The pulse-wave Doppler measurement is done with the ultrasound machine in standard zoom mode at a depth of 4 cm.

5.3.3	Image 3: Common Carotid Artery

Image 3 is a view of the distal 10-mm of the common carotid artery. Magnification is turned “ON”. The carotid bulb is displayed on the left side of the monitor (when facing the screen). If the bulb cannot be identified, but the tip of the flow divider can, this may substitute as the internal landmark on this view. After locating the tip of the flow divider and centering it on the screen, the probe is rotated into the lateral plane and moved downward 2 cm. The probe is then centered on the upper 1 cm of the common carotid. The probe should be angled to obtain the best view of the common carotid artery intima-media thickness (CCA IMT). In general the best angle to achieve the optimal CCA image is 45 degrees.

5.3.4	Images 4, 5 and 6: Internal or Bulb at the site of maximum wall thickness

Images 4, 5 and 6 are images of the internal artery or the bulb. Magnification is left “ON”. Images are centered on the site of maximum wall thickness in the internal carotid artery. The objective is to image the segment of the internal that contains the single largest wall abnormality, on either near or far wall, in any one view. The images should never be centered on the common carotid even if it is the site of maximum disease. The location of the site of maximum wall thickness is determined during the initial scan. It is not the site of average maximum wall thickness, but rather it is the single site of maximum wall thickness in any single view. The sonographer can use any available technique to make this decision including long axis or short axis views, color or pulse-wave Doppler.

5.3.5	The three internal carotid images are taken from three different scanning angles, the anterior-oblique, the lateral and the posterior-oblique. The point of maximal wall thickness should be centered in the middle of the screen for each image. The sonographer should adjust the probe to maximize the lesion and wall interfaces in each projection to not exaggerate the size of the focal plaque by scanning across the vessel on an oblique axis. The Ultrasound Reading Center readers are completely dependent upon the sonographer to provide a legitimate display of a plaque. The philosophy of the URC is that image drives angle rather than angle drives image.

5.3.6	The anterior-oblique, lateral, and posterior-oblique scanning angles are defined as follows:

	Anterior-oblique: the arch on the surface of the neck from the midline (trachea) to 55o to a line drawn from the mid-trachea to the center of the back of the neck.

	Lateral: the arch along the lateral surface of the neck, from 55o to the perpendicular to 100o (hence 45o). The sternomastoid muscle can be palpated beneath this portion of the skin’s surface.

	Posterior-oblique: the arch from 100o to the perpendicular to 145o. The probe almost always lies just behind the posterior margin of the sternomastoid muscle.

(Probe angle ranges are defined above to help guide the sonographer. The URC emphasizes optimal images over specific probe angle, however, we don’t want three identical views of the ICA.)

5.3.7	If the artery is normal, the study will be centered on the initial 10 mm of the internal carotid artery. The tip of the flow divider is used as the point that identifies the most caudal portion or inferior boundary of the internal carotid artery. It will not always be possible to identify the tip of the flow divider in the frozen image. The priority is to display the wall interfaces of the proximal internal artery, not the tip of the flow divider.

5.3.8	A cine loop of each of the four images described above, one of the common carotid and three of the internal carotid arteries, is captured by selecting the freeze button. The sonographer then cycles through the cine loop images to select the one that best displays the intimal walls. It is videotaped for five seconds. The cine loop is then set in motion and it is videotaped in its entirety.

6. 	Criteria for Satisfactory Images

6.1	The criteria for optimal B-mode ultrasound image of the carotid arteries is defined as the clear visualization on long axis views of arterial interfaces, internal arterial landmarks, and lesions.

6.2	Near wall - arterial wall nearest the probe

Adventitial - medial boundary

Intima  - lumen boundary

Far wall - arterial wall furthest from probe

Lumen - intimal boundary

Medial - adventitial boundary

6.3	The area of interest will be centered in the middle of the image. Align the probe to show as much of the vessel cephalad and caudad as possible. The sonographer should optimize the visualization of the interfaces by adjusting the gain settings, beam steering, and probe placement.

6.4	It is expected that the lumen of a good carotid ultrasound study image will contain a significant amount of artifact. To clearly visualize the intimal linings, the gain setting will need to be set considerably higher than it is for a typical clinical study.

6.5	Satisfactory images are properly magnified and the magnification level is the same for all longitudinal views. If a sonographer mistakenly uses an incorrect magnification level or changes it mid-scan, he/she should make a note of the problem in the “Comments” section of the videotape log sheet. Without a note from the sonographer the reader will assume the magnification is the same for all of the images. Thus, the image will be calibrated incorrectly, resulting in erroneous measurements.

7. 	Imaging Priorities

Some participants will not have vascular lesions or easily visible intimal linings. For these cases the sonographer will prioritize the far wall intimal linings over the near wall intimal linings. For the more difficult cases, where there is some disease, the top priority is to provide a quality image of disease. Disease located in the carotid bulb presents a special case. When a lesion is detected in the bulb and the internal carotid artery is normal, capture images to characterize the bulb disease. Instead of capturing an anterior, a lateral and a posterior view of the internal carotid artery, image the carotid bulb disease from these three perspectives.

7.1.  Lesion: If there is a focal lesion in either the distal 10 mm of the common or in the proximal 10 mm of the internal, focus on the lesion.

7.2.  Bulb lesion: If there is a lesion in the bulb and the internal is normal, focus on the bulb lesion.

7.3.  Far wall interfaces: If there is no disease, the clear visualization of the far wall is more important than that of the near wall

7.4.  Near wall interfaces

8. 	Sonographer Response to a Significant Stenosis – ALERT

8.1	Some participants will have significant carotid stenoses, which are discovered, perhaps for the first time, during this examination. An ALERT is defined as an 80% or greater stenosis in the common carotid, the bulb or internal carotid artery. The only criteria used to estimate stenosis is the peak systolic pulse-wave Doppler. An 80% or greater stenosis is indicated by a pulse-wave Doppler measurement of 250 cm/s. Imaging data should not be used in arriving at this conclusion; its role is limited to determining the site of the abnormality.

8.2	If a sonographer believes a significant vascular abnormality is present, he should double-check this finding by repeating the Doppler measurement. Under no circumstances should this impression be conveyed either directly or indirectly to the participant by the sonographer. The clinic coordinator should be told immediately after the participant has left the scanning area. An inquiry is thereby triggered at the field center regarding the presence of relevant symptoms in the participant. It will be determined whether he is under care for the vascular abnormality and if necessary appropriate referrals will be provided

8.3	The responsibility of the participant’s health care is completely with the field center. Whenever a participant presents with what the sonographer suspects is a problem he is to communicate it immediately to the field center medical personnel. Do not wait for confirmation from the Ultrasound Reading Center. The scan will not be reviewed until at least a week later. The readers and the project manager are not qualified to provide any sort of diagnostic report.

9. 	A Good Carotid Ultrasound Image

9.1	Wall boundaries can be demonstrated with high-resolution ultrasound imaging. They appear as two parallel echogenic lines separated by a hypoechoic space in longitudinal views of the carotid arteries. The artery walls are usually best observed in the common carotid artery where the vessel courses parallel to the skin surface and thus presents a target which is at a right angle to the ultrasound beam.

9.2	The first echo along the far wall is derived from the lumen-intima interface and the second, normally brighter, echo along the far wall originates from the media-adventitia interface. Between these interfaces lies the media, which appears as an echolucent zone. The distance between the first two lines corresponds to the combined thickness of the intima and media. Because of its collagen content the adventitia is quite echogenic and appears as a bright zone highlighted along its inner margin by the media. However the periadventia, depending on location, is composed of loose areolar tissue and in most instances is echolucent.

Common Carotid Image

Internal Carotid Image

9.3	It is more difficult to image the interfaces when the near and far walls of the vessels are curvi-linear and not at right angles to the ultrasound beam. In the carotid bulb, where the walls flare, only short wall segments may be seen on any single frame. This same phenomenon is observed in the proximal portion of the internal carotid artery when the walls are not parallel and hence sub-optimally visualized. Other causes for loss of wall interfaces that are not related to scanning technique are the presence of plaques and the presence of fat in the arterial wall. The interfaces can also be seen along the near wall but the lines may be disrupted and the echoes weaker because the ultrasound beam is passing from tissue to fluid. At times, it is impossible to maximize both the near wall and far wall interfaces on the same image. On such occasions the priorities are to first optimize lesion, second the far wall, and third the near wall. Lesion refers to the site of maximum single wall thickness, near or far wall.

9.4	The very earliest changes that can be seen in the arterial wall indicating the onset of atherosclerosis are fatty streaks. It is not clear that this finding can be detected by ultrasound. Fat is weakly echogenic and would not be expected to generate a signal. Picano et al. pointed out that the gaps seen in the lumen-intima interface, the "first interface" of the double line echo pattern along the far wall in a study of 75 fresh human aorta specimens might be a reflection of the presence of fatty plaques. However, such breaks may be present for multiple other reasons, including less than optimal scanning technique.

9.5	The first definite abnormality that can be detected by B-mode is intimal-medial thickening. As progressive development of atherosclerotic lesions occur, there is an increase in the error of estimating anatomic abnormalities because acoustical shadowing is often present with large lesions. Doppler, on the other hand, becomes progressively more accurate with increasing levels of disease. The Ultrasound Reading Center uses both grayscale imaging information and Doppler to compile a complete lesion profile for each carotid artery. It is important that both portions of the examination be done well.

II.	Overview of Carotid Ultrasound Analysis Protocol

1.	Preview and Digitization of a Carotid Ultrasound Study

1.1	The first thing the reader does when analyzing a carotid ultrasound scan is to read the sonographer’s comments on the log sheet and preview the videotape of the entire study. Through the preview process the reader learns of any unusual situations specific to the study, such as images videotaped out of the standard order. Carotid system structures are identified: the common carotid artery, the carotid bulb, the internal carotid artery, and if visible the external carotid artery. The site where the interfaces are most clearly imaged is identified in both the common carotid and non-diseased internal/bulb images. If there is disease in the internal/bulb, the site of the largest lesion is determined. In addition, the pulse wave Doppler measurements are assessed.

1.2	After previewing the videotape, the images are digitized. The computer-based ultrasound image analysis system combines the digitization and analysis functions. The program is designed to permit the user to immediately proceed with the analysis after digitizing a study.

2.	Calibration

The first step in carotid image analysis is calibration. It is necessary to calibrate each ultrasound image in order to convert the computer unit, pixels, to "real world" units, centimeters. To calibrate an image the reader positions the calibration tool such that it demarcates the 1-cm distance between two calibration marks. Readers assume that the magnification level does not change within a Carotid IMT scan.

3.	The Interfaces

3.1	The reader’s goal is to draw all six of the lines as shown on the each carotid artery image. The figure below demonstrates the anatomy of the carotid artery in the longitudinal perspective and the correspondence of the vessel layers to the line numbers used by the ultrasound image analysis program.

3.2	The correspondence between the line numbers and the interfaces are defined:

| Wall | Line Number |  | Interface |  |
| --- | --- | --- | --- | --- |
| Near Wall | Line 1 |  | Adventitia – Periadventitia |  |
|  |  | Line 2 |  | Media – Adventitia |
|  |  | Line 3 |  | Lumen – Intima |
| Far Wall |  | Line 4 |  | Lumen - Intima |
|  |  | Line 5 |  | Media – Adventitia |
|  |  | Line 6 |  | Adventitia – Periadventitia |

3.3	Measurements from carotid study images are made from the six lines drawn on the image. The lines are drawn with the stylus pen and are approximately 1-cm long. In the common carotid view, the measurements are taken in the 1-cm segment proximal to the bulb. The common carotid view is the only view from which measurements are made in the common carotid artery. It is a relatively simple task for a trained reader to draw the lines on a high quality image of a healthy artery. It is more difficult to analyze diseased arteries.

3.4	The measurements are that are made are of the near wall, far wall, vessel width and lumen. The measurement algorithm calculates the distance between pairs of lines and reports the minimum, maximum and average (and standard deviation) values. The vessel structures are defined by the lines pairs as follows:

| Vessel Structure | Line Pair |
| --- | --- |
| Near Wall | 2 – 3 |
| Far Wall | 4 – 5 |
| Vessel Width | 1 – 6 |
| Lumen | 3 – 4 |

4.	Priorities

4.1	The goal is to draw six lines on each carotid image. Due to varying degrees of image quality, which is highly dependent upon sonographer skill and subject anatomy, it will not always be possible to draw all six lines. It is, therefore, necessary to set priorities. The physics of ultrasound make the far wall is easier to image. The far wall interfaces will typically be more distinct than those of the near wall. Thus the far wall data set will be more complete and therefore be considered most important. Wall thickness is of greater importance than vessel thickness because this is a study of atherosclerosis.

4.2	The carotid ultrasound analysis priorities in order of importance are:

| Measurement | Line Pair |
| --- | --- |
| Far Wall | 4 – 5 |
| Near Wall | 2 – 3 |
| Vessel Width | 1 – 6 |

5.	Image Quality Scores

5.1	An image quality score is assigned to every analyzed image. This score is based on the quality of the image and the reader’s confidence that the lines drawn represent the true interfaces. Although the score is subjective, the following is a loose guideline for scoring the common and non-diseased internal carotid arteries. Note that the criterion used to assign a quality score to a diseased vessel differs slightly from that of a healthy vessel

5.2	Common & Healthy Internal Arteries

	Unacceptable: only one or no lines drawn.

	Poor: only two lines, one pair, drawn, giving a lumen measurement (i.e.. 3&4, or 2&5) or vessel width (i.e.. 1&6)

	Acceptable: only two or three lines on either near or far wall with an opposing line (lumen measurement).

	Very Good: all six lines drawn, all measurements made.

	Excellent: all six lines clearly visualized and easily drawn, with good gain and positioning.

5.3	Atherosclerotic Carotid Arteries

	Unacceptable: only one or no lines drawn.

	Poor: only two lines, one pair, drawn, yielding a lumen measurement (i.e.. 3&4, or 2&5) or vessel width (i.e.. 1&6) or can draw lines on wall opposing the lesion but no lesion measurement can be made.

	Acceptable: four to five lines drawn, lesion traceable but not clearly defined, with at least one posing line.

	Very good: all six lines drawn, lesion clearly defined, with opposing line.

	Excellent: lesion clearly defined, with opposing wall measurement, lines are easily drawn and the image has good gain and positioning.

6.	Subjective Assessment

In addition to obtaining measurements of wall thickness, vessel width and lumen the reader characterizes the internal/bulb disease. Lesions in the common carotid artery are not characterized. Assessments are made of the surface, morphology, percent stenosis, location and density. Each characteristic and the categories the Ultrasound Reading Center uses to qualify them are briefly described.

6.1	Surface

The surface of the artery or lesion is categorized as one of the following:

	Smooth - no lesion is present in the artery and the lumen-intima surfaces (lines 3 & 4) are free of irregularities. If there is a lesion present then this category is applicable when the lumen-intima surface (lines 3 & or 4) of the lesion is smooth and free of irregularities.

	Mildly irregular - the lesion has minute surface irregularities.

	Markedly irregular - the lesion has noticeable surface irregularities. This can be best described by saying that the surface is “bumpy” but without a prominent pit or crater.

	Ulcerated - the lesion surface has prominent ulceration, pit(s)or crater(s) that is at least 1mm in depth and is at least 1mm in width. The “back wall” of the lesion, defined as the media-adventitia interface, can be clearly seen at the point of ulceration.

	Can’t tell - the surface characteristics of the artery/plaque cannot be determined due to insufficient information (images are unclear or missing).

6.2	Morphology

The morphology, the form and structural characteristics, of the lesion are categorized as:

	No lesion - no lesion present in either the internal or bulb

	Homogeneous - a lesion that demonstrates an echogenicity that does not vary across its width and thickness.

	Heterogeneous - a lesion with have a mixture of lucent and echogenic zones.

	Can’t tell - the morphology of the plaque cannot be determined due to insufficient information (images are unclear or missing).

6.3	Percent Stenosis

6.3.1	All available information is used to determine the percent stenosis. Information sources include digitized images, cine loops and Doppler values. The degree of stenosis is crudely estimated from the images and cine loop as a ratio of artery width. The Doppler value (velocity) which is known to be a more reliable measure is incorporated to refine the estimate.

6.3.2	It may not be evident from the images the images that there is a 100% occlusion. In some cases, there may not be a Doppler value because the sonographer was unable to produce a Doppler signal. When Doppler values are nonexistent and the images are unclear, the reader must rely on comments from the sonographer to determine the degree of stenosis.

6.3.3	In all instances where the PW Doppler is greater than 1.5 m/s, it is used to categorize the percent stenosis according to the following table.

| % Stenosis | Doppler Values |
| --- | --- |
| 0-49% | <1.5 m/sec |
| 50-74% | > 1.5 m/sec but < 2.5 m/sec |
| 75% -99% | > 2.5 m/sec |
| 100% (occlusion) | 0.0 m/sec |

6.3.4	The categories from which the percent stenosis must be chosen are more specific:

	Normal - no focal lesion present in either the internal or bulb.

	1–24% - Doppler values that are <1.5 m/sec and reader’s estimate based on the three images of the internal showing that the lesion occupies 1-24% of the vessel width.

	24–49% - Doppler values that are <1.5 m/sec and reader’s estimate based on the three images of the internal showing that the lesion occupies 25-49% of the vessel width.

	50–74% - Doppler values that are > 1.5 m/sec but <2.5 m/sec.

	75–99% - Doppler values that are > 2.5 m/sec.

	100% - Doppler values that are 0.0 m/sec.

	Bad image/Nl (normal) Doppler - the images are so unclear that % stenosis can’t be estimated by looking at them but Doppler values fall within normal range, < 1.5 m/sec.

	Can’t tell – the percent stenosis cannot be determined due to insufficient information (images are unclear or missing).

6.4	Location

The plaque location describes the site of the lesion. The categories are:

	No lesion - no lesion is present in either the internal or bulb

	Can’t tell - cannot determine the location of the plaque cannot be determined due to insufficient information (images are unclear or missing).

	Internal - the lesion is located in the internal carotid

	Bulb - the lesion is located in the bulb

	Can’t tell - cannot determine the location of the plaque cannot be determined due to insufficient information (images are unclear or missing).

6.5	Density

The lesion density qualifies the compactness of the plaque. The categories are:

	No lesion - no lesion is present in either the internal or bulb

	Hypodense - a plaque that appears darker than the surrounding tissue.

	Isodense - a plaque that has same level of brightness as the surrounding tissue.

	Hyperdense - the plaque appears brighter than the surrounding tissue and some of the tissue beneath the plaque is shadowed.

	Calcified - the plaque appears much brighter than the surrounding tissue and all of the tissue beneath the plaque is shadowed.

	Can’t tell - cannot determine the density of the plaque cannot be determined due to insufficient information (images are unclear or missing).

III.	Data Transmission

1.	Videotape

1.1	Carotid ultrasound images are transmitted to the Ultrasound Reading Center on videotape. The ultrasound machines are equipped with Super VHS videocassette recorders. Super VHS tape must be used. It can be purchased at Radio Shack or through the field center’s vascular lab or radiology department. Standard, non-Super VHS, tape is not acceptable.

1.2	When the image displayed on the ultrasound machine monitor is optimal, “Freeze It”. Page through the cine loop of saved images and select the image that best demonstrates the intimal linings. Videotape five seconds of the frozen image, start the cine loop on loop preview and videotape it in its entirety. It is important to wait for the mechanical delay of the VCR mechanism to engage when timing recordings. A few seconds of over-recording is infinitely more desirable than a non-recording. If at any time during an examination a better image than that previously recorded is obtained, do not erase or tape over the previous recording. Add the new images to the videotape and make a note in the Comments section of the log sheet to inform the reader.

1.3	At the Ultrasound Reading Center the readers digitize (computerize) and analyze the images from the carotid ultrasound videotapes. The computer-based system to accomplish these tasks was developed by the Ultrasound Reading Center.

2.	Labels and Log Sheets

2.1	At the training session the sonographers will be given sets of labels and master copies of a log sheet. These items make it possible to track and organize the numerous videotapes that arrive at the Ultrasound Reading Center. It is of utmost importance that the sonographers carefully label the videotapes and the corresponding log sheets. The labels contain the study name, field center name, exam number and tape number. See the label illustrated in the upper right corner of the sample log sheet below. The log sheet enables the reader to quickly locate each study on the individual videotapes and to obtain pertinent information regarding particular cases.

2.2	The sonographer completes the appropriate sections of the log sheet for each subject he scans. The necessary items are relatively obvious.

	Scan Date – the date of the carotid scan

	Subject ID – the ID number assigned to the participant (no names please)

	Plaque – if there is a plaque(s) circle the location(s). RB = right bulb, RICA = right internal carotid, LB = left bulb, LICA = left internal carotid

	VCR Start – the VCR counter/timer that indicates where the particular study is located on the tape

	Sonographer ID – the MESA Technician ID assigned to the sonographer

	Comments – the place for the sonographer to tell the reader about the case

	Sample of Typical Videotape Log Sheet Entries

|  | Scan Date | Subject ID | Plaque  | VCR Start Time | Sonog. ID | Comments |
| --- | --- | --- | --- | --- | --- | --- |
| 1. | 8/2/2000 | 4125993 | RB RICA LB LICA | 0:00 | 456 | Good study, easy subject |
| 2. | 8/2/2000 | 4328713 | RB RICA LB LICA | 5:34 | 413 | Lesion in left bulb |

2.3	The value of sonographer - reader communication cannot be overemphasized. Use the Comment section of the log sheet to communicate any abnormalities, problems or difficulties encountered. For example, if a scan is canceled or if there are images missing or out of order, write it down. The readers base all of their decisions on the information the sonographer provides. It is not easy for a reader who is not medically trained, to distinguish plaque from artifact. Efforts by a sonographer to assist the reader in the analysis process are greatly appreciated.

3.	Shipping Videotapes and Log Sheets

3.1	The videotapes are to be sent monthly to the Ultrasound Reading Center. The package should include the videotape, log sheet and a hard copy packing slip. A copy of the log sheet should be made prior to packaging it with the videotape. This will enable the sonographer to maintain a record of his scanned subjects. Mailing labels will be provided at the training session, the delivery address is:

Daniel H O'Leary, MD

Ultrasound Reading Center

25 Stuart Street, 2nd Floor

Boston, MA 02116

3.2	The Ultrasound Reading Center insists that a delivery service with tracking capabilities be used. This will reduce the possibility of lost data. United Parcel Service ground delivery has proven to be both satisfactory and economical.

4.	Data Transmission to the Coordinating Center

The carotid ultrasound analysis program output is an MS Access database that contains the calculated measurements and subjective scores for each analyzed participant study. The calculated measurements include values for the far wall, near wall, vessel width and lumen. The subjective scores that characterize the surface of the artery or lesion include assessments of the morphology, the percent stenosis, the location of the lesion, and the morphology. Also included with each participant data record are items such as sonographer ID, scan date, reader ID, read date and pulse-wave Doppler measurements. The data files will be appropriately formatted and sent on a regular basis to the Coordinating Center on either floppy disk or via e-mail.

IV.	Sonographer Quality Control

Following initial certification sonographers will be evaluated continuously to monitor and assure quality control. Since reliable data is the essential element in studies that seek to monitor arterial change over time, great effort must be expended upon all aspects of quality control. The Ultrasound Reading Center has found it best to minimize the number of sonographers and to maintain the same reader thereby limiting problems prospectively rather than retrospectively adjusting data.

1.	Field Center Sonographers

1.1	If possible, two sonographers should perform all examinations at each field center on the same ultrasound machine. Ideally, these individuals would be an experienced vascular sonographer, preferably RVT certified or equivalent. If not the sonographers should have experience in an ICAVL accredited laboratory. The sonographers do not necessarily have to have experience performing research protocols. The Ultrasound Reading Center Carotid IMT protocol is such that it can be quickly taught to someone who routinely performs high quality clinical ultrasound examinations.

1.2	The Ultrasound Reading Center assumes that at most of the six field centers there will be two sonographers performing all of the carotid studies. Together with the Coordinating Center the Ultrasound Reading Center will assess inter-sonographer variation. Intra-sonographer variability measurements, which will involve the replicate scanning of approximately twenty subjects by the same sonographer on two different occasions, are highly recommended.

2.	Ultrasound Reading Center – Sonographer Communication

2.1	Experience with prior studies demonstrated the need for rapid and possibly frequent communication between the Ultrasound Reading Center and the sonographers. The project manager and readers will periodically call each of the field centers. Problems encountered by the sonographers and sonographer performance will be discussed and the importance of sonographer compliance with the protocol will be stressed. Examples of the kind of problems that required immediate attention in the past include the apparent non-recording of images on videotape and discordance between Doppler and imaging data. Sonographers are asked to keep notes when problems are recognized so that they may provide direction to the Ultrasound Reading Center when necessary. Sonographers are encouraged to call the project manager to discuss problems or questions as they arise. If necessary, a member of the Ultrasound Reading Center staff will visit field centers. The execution of the carotid protocol will be observed and equipment function evaluated. The purpose of the visits is to be a direct assistance to the sonographers and the field centers.

2.2	If a sonographer submits unacceptable or un-analyzable studies, the reader will alert the project manager. An unacceptable study is one from which very little or no data can be obtained. Unusable studies can be the result of poor images or failure to follow the protocol (incorrect images). Poor performance by a sonographer will result in a telephone call to the sonographer and the clinic coordinator from the project manager. Severe cases will warrant a memo to the field center and Ultrasound Reading Center Director documenting the persistent poor quality.

2.3	Throughout the exam period, the sonographer will receive sets of Reader Comments.  The Reader Comments are written by the reader at the time of study analysis. Typically, the comments contain constructive criticisms and suggestions for improving the study images, i.e., gain adjustments, check the Doppler gate angle or position. More often than not, the Reader Comments include such positive feedback as “Great images, thanks”.

3.	Ultrasound Reading Center Report on Sonographer Performance

3.1	In addition to the Reader Comments, sonographers will receive feedback from the Ultrasound Reading Center in the form of a sonographer performance quality control report. The maintenance of high performance requires immediate and direct feedback. The reports, which will be published bi-monthly, quantitatively track each sonographer’s performance. Copies of the reports will be distributed to the field centers, the Ultrasound Reading Center PI and the coordinating center. Cumulative summaries will be published periodically to serve as an ongoing benchmark against which to judge individual results.

3.2	The report will include the number of studies, average image quality scores and percentages of possible lines drawn and Doppler velocity measurements tabulated for all of the field centers, and then by each individual field center, by sonographer and reader. A one page summary report that highlights the overall study results as well as the best and the worst field center and sonographers will accompany each report. Consistently poor performance by a sonographer will initiate a call of encouragement and investigation by the project manager. Continued poor performance will result in the sonographer being re-trained or replaced.

4.	Performance Standards During Examination Period

A baseline quality control score, derived from the score assigned to each image and the number of lines drawn for each image, will be established after analysis of pilot and pre-pilot data, below which sonographer performance will be considered unacceptable if maintained for three consecutive weeks. The readers will be monitored to ensure uniform scoring throughout the exam period. Additional training may be required of sonographers and/or readers in the event of any significant temporal drift in accuracy or precision of sonography study performance and/or reading. In terms of the analysis of the actual measurements produced, we will subscribe to those suggested by echocardiography - namely, a goal of mean variability of 10%, with a cut-off point of 15% acceptable variability.

5.	Replacement or Retraining of a Sonographer

If a sonographer resigns or is replaced due to poor performance, a replacement must be trained and certified prior to performing scans on research subjects. The Ultrasound Reading Center must conduct all sonographer training. A new sonographer can be introduced to the protocol by a certified sonographer, but the new sonographer will not be certified to conduct participant exams until he attends an Ultrasound Reading Center training session. Contact the Ultrasound Reading Center Project Manager immediately when the need to train a new sonographer becomes apparent. Due to the Ultrasound Reading Center’s involvement in numerous national multi-center research studies, it may be possible to identify a local sonographer who is certified to execute the carotid ultrasound protocol to assist with the training.

6.	Within Sonographer Variability Assessment

Intra-sonographer variability during the examination period will be assessed from the data obtained by repeating complete carotid ultrasound studies on a number of subjects. Each sonographer will complete two scans on a subset of approximately twenty participants distributed throughout the exam time period. The videotapes will be sent to the Ultrasound Reading Center for analysis coded with random number assigned by the Coordinating Center. The reader will be blinded – he will not know that a study he is analyzing is a standard or a sonographer intra-variability quality control study. Subjects for the replicate studies will be recruited during the exit interview. They will not be recruited by the sonographer. The repeat scan should be scheduled to occur no sooner than one week and no later than one month after the initial scan. On the return visits, each volunteer will be scanned once by each sonographer.

7.	Achieving and Maintaining Quality Ultrasound Data

7.1	Quality carotid ultrasound data begins with the sonographer. The sonographer produces the images from which the readers extract the date. A poor set of images results in a poor and sometimes incomplete set of data. The readers are entirely dependent upon the sonographers to provide clear images and to explain any conditions, anatomical or procedural that vary from the norm.

7.2	To maintain high quality ultrasound data:

7.2.1	Acquire nothing but clean images that clearly display the intimal linings and the lesions. The intimal linings and the lesions are the top priority, they are what is measured by the reader. It is expected that in order to clearly display the intimal linings the lumen will contain more artifact than the typical clinical study. Review Section 2.4, “Priorities.”

7.2.2	Videotape the images in the order stated by the protocol. A study that is videotaped out of order is not as easy for the readers to analyze.

	Right transverse sweep of common and internal

	Right pulse wave Doppler measurement

	Right common carotid

	Right internal/bulb anterior oblique

	Right internal/bulb lateral oblique

	Right posterior/bulb oblique

	Left transverse sweep of common and internal

	Left pulse wave Doppler measurement

	Left common carotid

	Left internal/bulb anterior oblique

	Left internal/bulb lateral oblique

	Left internal/bulb posterior oblique

7.2.3	Annotate the images with the following information:

	MESA

	Subject ID

	Subject Acrostic

	Sonographer ID

	Image acronyms: RPW, RCC, RAO, RLO, RPO, LPW, LCC, LAO, LLO, LPO

7.2.4	Magnify the views of the CCA and ICA. Keep the magnification level constant throughout the scan.

7.2.5	Annotate the images and use arrows to distinguish the vessels when the images are atypical. Don’t force the reader to guess which vessel is the internal and which is the external carotid.

7.2.6	Use the Comment section of the log sheet to communicate any abnormalities, problems or difficulties encountered. For example, if a scan is canceled or if there are images missing or out of order, write it down. The importance of this communication cannot be stressed enough. The readers are not medically trained and must base their decisions on the information the sonographer provides.

7.2.7	Videotape the frozen images for 5-10 seconds. Any less than 5 seconds makes the image difficult to capture and any longer than 10 seconds reduces the reader’s efficiency.

7.2.8	Videotape the complete cine loops. Cine loops enable the reader to discriminate between true anatomy and artifact.

7.2.9	Take the time to verify subject ID’s. Check the ID typed into the ultrasound machine and written on the log sheet. If an incorrect ID is videotaped because the wrong ID was typed into the ultrasound machine, tell the Ultrasound Reading Center about it with a note in the Comments section of the log sheet. After the videotape leaves the field center sorting out erroneous ID’s can be a very difficult, if not impossible task. Watch for transpositions.

3.9  	Carotid MRI – non-contrast and Gadolinium-Enhanced

There will be no Carotid MRI performed for Exam 3.  All participants completed the MRI as of Exam 2.

3.10 		Computed Tomography (CT)

I. 	Background and Goals

1.	The Multi-Ethnic Study of Atherosclerosis (MESA) is studying risk factors and measures of cardiovascular disease that relate to progression of subclinical to clinical disease.  An integral part of this study is the measurement of coronary artery calcium using either electron beam computed tomography (EBCT) or helical computed tomography (HCT).   Coronary calcium will be assessed in relation to the risk of future cardiac events, and from repeated scans in selected individuals, the progression of coronary calcium will be related to baseline risk factors and risk of future events.

2.	Of the participants that have been recruited in the MESA, a randomly selected sample of about half of participants will receive repeated scans for coronary calcium during Exam 2.

3.	Goals

3.1	To examine the relation of baseline measures of coronary calcium quantity (score and volume measures) to future risk of cardiac events

3.2	To examine the relation of baseline measures of coronary calcium to other risk factors and measures of subclinical disease (e.g., MRI and ultrasound measures).

3.3	To examine the relation of progression of coronary calcium to the risk of future cardiac events, as well as to risk factors at baseline and progression of risk factors.

3.4	To examine the above goals by ethnic and gender strata, when possible.

II.	Qualifications of Personnel

1.	As with Exam 1, Field Center personnel will be responsible only for escorting participants to and from the CT scanning site and ensuring that the ‘CT Completion Form’ is completed by the scanning technologist.

2.	Each Field Center has a designated radiologist or cardiologist who is responsible for the performance of the CT examinations at his/her MESA Field Center.   This physician’s role is similar to that in Exam 1.

3.	Field Center technologists should have appropriate knowledge of cross-sectional anatomy, physiology, and pathology related to the heart.

4.	Each technologist involved in the study should also have a complete understanding of this protocol, be experienced at providing breath-holding instruction, ECG gating, operation of the EBCT or HCT equipment.

5. 	To ensure quality control, each Field Center should have designated CT technicians who will perform the MESA examinations. Technologists must be certified as RTs in their state. Technologist certification needs to be fulfilled as in Exam 1.

III.	Participant Scheduling

1.	Procedure is unchanged from Baseline Exam.  Refer to the “Exam 1 MOP” for further details.

2.	A Field Center interviewer/scheduler will be responsible for explaining and obtaining consent for the CT examination.  Participants will be scheduled for a certain date and time, and transportation arrangements, if necessary, will be arranged by the Field Center interviewer/scheduler.  A confirmation appointment letter will be sent, providing the time, date, and directions to the scanning center, and describing the procedure.

3.	The scanning procedure will require approximately 20 minutes of the subject’s time.

IV.	Completing the CT Completion Form

- As with Exam 1, the CT Completion form must be filled out by either the CT technologist or the attending Field Center staff member.  If the Field Center staff member completes the form, he/she will need to get information from the CT technologist before leaving the scan area.  The form will be taken back to the clinic and scanned into the computer along with the other forms. The information on the form will be forwarded to the CT reading center for image tracking purposes.

Note:  Record current clinic weight—not self-reported weight—in the “clinic weight” field and enter clinic weight into the CT machine prior to performing the scan.

  - Question 1:  For women 45-55:  Results of pregnancy test.

- Bubble in the appropriate choice: “Positive”, “Negative”,  “Not Done”.  If not done, state reason in space provided.

1.2	Question 2:  Results of CT scan.  Bubble in the appropriate choice below:

	Complete:  Select this response if the scan was complete or partially complete and a file containing the scan data will be sent to the CT Reading Center.  Then go to question #3.

	Incomplete:  Select this response if the scan was only partially completed due to a technical or participant problem.

	If Incomplete is selected, indicate whether a file containing this exam information will be sent to the Reading Center.

	Not Done:  Select this response if the scan was not even started, due to a technical or participant problem.

	If Incomplete or Not Done is selected, indicate the reason the scan was not completed.

1.3	Question 3 is a comments section.  Note any comment you would like to RC to consider in evaluating this CT scan.

3.11   CT Calcium Scoring  and Diameter of the Aorta

I. Description of Boxing

In the boxing window, you will view consecutive images of the abdominal aorta. You will then enclose aorta in the white box. This procedure will produce the boxed view of the aorta you will be using during the Tracing and Scoring procedure.

Another purpose of the boxing stage is to properly identify the hydroxy-apatite phantom regions.

To first get started you will have to launch the boxing program. On the subject chooser highlight a patient study by first clicking on it. Then move the mouse over the first icon labeled “ Take Heart”.

When the boxing window pops up, it will cycle through all the slices of the patient’s images. When this happen you will keep an eye on the aorta to make sure that the aorta is well enclose within the white box. To check the slices again after the initial cycle, right-click anywhere on the boxing window (to bring up the menu) and choose “ Cycle Through”.

II. Boxing the Aorta

This serves to isolate the image in the boxed portion for the next process.  After the window finishes cycling through the images, there should be white boxes surrounding the image of the aorta. Resize the box by left clicking on either corner and stretching the white outline so as to completely encompass the aorta.

Once again you will have to check if  the  aorta has been included in the boxed region. Do this by right clicking to being up the menu and then choosing “Cycle Through “. Once you are sure the entire aorta is boxed bring up the menu once again and click “Done: process”. A small window will pop up on the top left corner of the screen, choose “begin tracing“ to launch the second part of the Scoring process.

III. AORTIC DIAMETER

- Determine at which level the aortic bifurcation occurs. This will be the most caudad level for which the aorta is a single vessel.  Then determine which levels are 8 cm and 5 cm cephalad to the bifurcation.  On EBT scans this will be 17 and 9 slice levels from the bifurcation.

- Go to the bifurcation level.  Right click and select "Draw Polygon". Using the left mouse button to drag a circle around the aorta at the  three levels. First, 8 cm (or 13 slices, approx. 78 mm for EBT) proximal to bifurcation, 2nd   5 cm (or 8 slices, approx. 48mm for EBT ) proximal to bifurcation) and 3rd,   bifurcation and hitting ctrl+c to save the area. Then repeat this step for the two levels that are 8 and then 5 cm cepalad to the bifurcation.

- Remember that the order for measuring the size of the aorta is 8 cm proximal to the bifurcation, 5 cm above the bifurcation.

- If a mistake is made, note which measurement (the first draw, second, third etc...). After the aorta has been measured on all three levels and the tracing has been completed as normal, continue through the scoring.

- After scoring, if there are more or less than three measurements of the aortic size,  a window will appear with simple onscreen directions to select the three to be scored or reminding the reader to repeat the measurements ( if there were less than three).

IV. TRACING

##### Launching the Tracing Program

There are two ways to launch the Tracing  program, by clicking on the Trace button or by  finishing the  Boxing process as described in the  previous section.

The purpose of the tracing program is to actually to trace the path of the aorta through the abdomen so that the Scoring Software will know where to look for the lesions. Basically you will scroll through all the slices and identify where  the appropriate vessels are. Navigating through the slices is simple . First you must left click the tracing window. After you may move forward through the slices by pushing the ‘down’ arrow key on the keyboard. To go on reverse, use the ‘up’ arrow key.

- Start tracing at the very first slice using the LM options (red xs) using the method described below.  When you reach the level that is 8 cm cephalad to the bifurcation go to the next step.

- Label the upper right hand quadrant of that image level (8 cm cephalad to bifurcation) as LAD and then trace the LAD Br1 X’s in a clockwise direction around the aorta being sure to put enough X’s so that the resulting ROIs completely encompass the wall of the aorta.

### LAD

LAD(Br1)

In this way you can include the walls of even the largest aortas.

3.	On the next slice put the LAD X exactly below the one on the slice before and start tracing with the LAD Br1, arrows in the opposite direction (counterclockwise) as follows:

LAD

LAD(BR1)

4.	Continue in this way until you have completed all the slices up until and including the slice where the aorta flattens out (right before it becomes two vessels).

5.	At the first level where there are two separate iliac arteries, use the RCA arrows for the right iliac and the CX arrows for the left iliac. Use the RCA and CX (Br1) arrows to surround the iliac in the same way that you used the LAD (Br1) arrows to surround the aorta.  Trace thusly to the last slice.

In the end you should have a trace that looks like this:

Checking the Tracings

After you have labeled the vessels in the last slice where vessels appear, you should check your work by right clicking anywhere on the image and choosing: Done: process” from the pop up menu. When you do this you will see the images re appear consecutively with the circular regions surrounding the points, where you placed the X’s. Be sure that the entire aorta and the bifurcation are included.

Finish Tracing

If the regions are placed correctly, choose either : Save and Score” from the resulting options, to proceed to the Scoring program or “Exit” to end tracing and return to the Subject Chooser. If the regions are not correct, press “ edit” and add X ‘s as needed.

V.  SCORING

Launching The Scoring Program

Launch the Scoring program by either highlighting a Traced study in the Subject Choose and clicking on the Scoring Icon, or by choosing “Save and Score” when finishing the Tracing Program, see above. Proceed exactly as you would for  ,scoring the coronaries.

### VI.. QUALITY  ASSESSMENT  TOOL

The QA menu , has seven categories that you must assess. These are Motion, Streak, Phantom, Mis-registration, Noise, Arteries and Coverage.

Choose your assessment  by using the following guidelines.

Motion-Use this if there is a  motion artifact in the image

Streak- Choose ‘Unacceptable ‘ if streak artifact overlaps any  of ht e three regions selected.. Otherwise choose excellent.

Phantom-Choose ‘Unacceptable ‘ if insufficient assessment of the phantom occurs in any image. Otherwise choose ‘Excellent’.

Mis-registration- Choose ‘Average’ if the images in the projection window do not show smooth borders . Otherwise choose ‘Excellent’.

Noise- Choose ‘ Average’ if t image contains excessive noise such as the program has mistaken noise for coronary calcium. . Otherwise choose ‘Excellent’.

Arteries-Choose ‘ Unacceptable if it does not reach the bifurcation. Otherwise choose  ‘Excellent’.

Coverage- Choose ‘Unacceptable’  if it does not reach 8 cm proximal to the bifurcation. Otherwise choose ‘Excellent’

Once finished, click the “ Save” button and Exit.

V.	Reading Center Alerts

The Reading Center will review all scans within two weeks and will indicate any alerts.  A cardiologist or radiologist investigator will review all scans and will identify eventual alerts.  The investigator will telephone the CT investigator at the field site and the coordinating center and will write a letter (see CT Manual of Operations) with copies to both. Pathological findings constituting alerts will include pulmonary infiltrates, pericardial or pleural effusions, tumors, dilatation of the aorta greater than 4 cm and pathological rib fractures.

SECTION 4:  ALERTS

I.	Purpose

The purpose of defining medical alerts is to make sure that the participant and his/her physician are aware of any significant medical findings that arise as a result of the MESA clinic exam.

II.	Definitions

1.	Alert:  Any of the medical findings, listed in Table 1, that may have adverse health consequences to the participant if untreated.

2.	Immediate Referrals:  Medical emergencies, which require immediate notification of both the participant and his/her primary physician.  Participants receiving immediate referrals should be considered as those who would go directly from the Field Center clinic to their physician or hospital.  Immediate notification of the participant should occur during the clinic visit.  Immediate notification of the participant’s physician should be accomplished by telephone, to be completed before the participant leaves the clinic.  A follow-up letter documenting information discussed by phone should also be sent to the participant’s physician for findings included in the initial report, or immediately upon receipt from the central agency for findings on the later reports.

- Urgent Referrals:  Urgent referrals are made for abnormalities detected, which require medical attention, but not on an emergency basis.  Urgent notification of the participant should occur before the participant leaves the clinic (for findings included in the initial participant report), or immediately upon receipt from the central laboratory (for findings on later reports).  Urgent notification of the participant’s physician should be sent within the week.

III.	METHODS

1.	General Instructions

Whenever one of the alerts listed above is identified for a participant, the actions defined under Immediate or Urgent Referrals, above, must be completed.

Table 1. MESA Alerts and Alert Levels

| # Finding | # Alert Level |
| --- | --- |
| Systolic BP > 210 | Immediate |
| Diastolic BP > 120 | Immediate |
| 180 < Systolic BP < 211 | Urgent |
| 110 < Diastolic BP < 121 | Urgent |
| Total cholesterol > 360 mg/dL | Urgent |
| Triglyceride > 1000 mg/dL | Urgent |
| HDL cholesterol < 20 mg/dL | Urgent |
| Calculated LDL cholesterol > 260 mg/dL | Urgent |
| Fasting glucose < 50 mg/dL or > 400 mg/dL | Urgent |
| Creatinine > 2.0 mg/dL | Urgent |
| Any abnormalities identified by CT tech | Urgent |
| Any abnormalities identified by MRI tech | Urgent |
| Any abnormalities identified by MESA Eye readers | Urgent |
| CT RC Alert | Urgent |
| MRI RC Alert | Urgent |
| MESA Eye RC Alert | Urgent |
| Ultrasound RC Alert | Urgent |

SECTION 5:  REPORTING PARTICIPANTS’ RESULTS

i.	mETHODS

1.	Reporting Participants’ Results

Participants receive three results reports.

- The lab results report contains the results of the blood values. This letter informs the participant if any of the results are abnormal. In addition, a letter is sent to the participant’s physician, along with a page of reference ranges and the participant’s lab results.

- The “high-tech” results report contains the results of the MRI, and CT scans and informs the participant of any abnormal finding. A letter containing a more detailed report of the scan results is also sent to the physician.

- The MESA Eye results report contains the results of pathology found in the retinal photography.

	All results letters should be carefully reviewed before they are sent to participants and physicians.

	The MRI section should be reviewed to see if further information or clarification is needed.  Findings, whether they are alerts or routine findings, will be inserted verbatim into the physician letters from the database descriptions that have been entered and transmitted by the MRI readers.  These may be very brief or incomplete sentences and Field Center staff may wish to enhance them for improved clarification.  If the finding is an alert, there will be a letter from the Reading Center to assist in this process.  The participant letters will include a description of the finding in layman’s terms.  However, if the finding is one that does not fit into the specific alerts, but is recorded in “Other Alert”, the description in the participant letter will be inserted from the database just as in the physician report.  In this case, medical personnel may wish to edit the participant letter using less technical terms for the finding reported.

- MESA Vision refraction will give participants a computer generated vision refraction report at exit time when the participant completes the exam.

SECTION 6:  DATA MANAGEMENT

6.1 Overview

The MESA data entry and management system has several components many of which are the same as for Exam 1.  This section will provide a highlight of some of the component.

There are four main icons on your desktop that will be used for accessing the MESA data system:

TELEform Reader

TELEform Verifier

MESA Data Management

MESA Events

Click on the TELEform Reader icon when you want to scan forms into the database.

Click on the TELEform Verifier icon when you want to verify previously scanned forms.

The “MESA – Data Management” icon contains all components of the MESA data tracking system.  In this database, each record represents an individual person.  Contained in this system are procedures for printing forms, entering recruitment status, checking information about individuals, and creating reports.  This system also contains procedures for transmitting data to the Coordinating Center and for performing data backups.

6.2 General Instructions for Completing Paper Forms that will be scanned.

Scanned forms are a fast and efficient way to enter data into a computer system, but it is important to complete the forms in a careful manner, with an understanding of how the scanning system operates.  This will ensure that the forms are correctly interpreted by the TELEform system.

TELEform’s “User Guide” recommends a fine felt tip pen.  It works best because it leaves solid line segments.  Some ballpoint pens create broken line segments that introduce greater possibility of recognition error.  A mechanical pencil using 2B or softer lead will also work; though it is very important make sure enough pressure is applied using one so that the marks are solid and dark.

Preferred	Acceptable but not recommended

A majority of the questions on the paper forms have option bubbles as valid responses (i.e. 0-No, 1-Yes, 9-Unknown).  Even though TELEform will usually recognize a partially filled in bubble (see above), completely filling it in is strongly recommended because doing so greatly reduces the chance that TELEform will miss an option that was selected.  If a question with bubble options is completed when it should have been left blank, the reviewer should mark through multiple bubbles for that question. This will cause the question to come up for verification and allow the correction to be made. If the reviewer just crosses out the incorrectly marked bubble, the verifier will still interpret that bubble as the correct response.

In any response area that accepts numeric or text entry, each number or letter should be printed neatly and should avoid touching the edges of the response boxes or comb-style fields (see above).  If a character is touching an edge, TELEform should still be able to recognize it, but in some cases, TELEform will not correctly recognize the character and will require that the operator verify the response.  All alphabetic characters should be printed in UPPERCASE CAPITAL LETTERS. See Appendix I, “Handwriting Examples for Teleform,” for examples of specific handwriting problems and how they were interpreted by Teleform.  Also included are examples of good handwriting of numbers and letters.

All dates should be entered in the format MM/DD/YYYY (note the four digit year) with no missing month, day, or year portion (“12/  /2002” would not be acceptable), and all times should be entered in the format HH:MM A/P with no missing hour, minute, or time of day portion.  Refer to Figure 2.2 for examples of valid date and time responses.

Responses entered using a mechanical pencil (with 2B or softer lead), can be changed by thoroughly erasing the prior response and then entering in the correct response.  If you make a mistake when writing in a response with a pen, cross out the incorrect response and enter the correct one, though this will force the TELEform operator to correct the response during verification.

Don’t use correction fluids like “Liquid Paper” because using such products may potentially dirty the roller on the document feeder or contact glass on the scanner.

6.3 Entering and Verifying Data

6.3.1	Using The Scanner And TELEform Reader To Enter Data

Make sure the scanner is on before turning the computer on.  This is necessary because the scanner is a SCSI (Small Computer System Interface) device and SCSI devices are not recognized if they are disconnected or off during the initial boot-up phase.  Once the computer is on, you can turn off the scanner and turn it back on later, but if the computer is ever turned off and back on (or re-booted), make sure the scanner is on first.

Procedures for using the scanner are the same as for Exam 1.  Refer to the “Baseline Exam MOP” for further details.

Important:

Each page has corner marks that must be unobstructed for TELEform to properly interpret the page; TELEform uses these marks to determine the relative locations of the data-entry areas.  An identification number that appears at the lower right hand corner of the form also needs to be visible.  Check all the pages to make sure there are no folded corners or extraneous marks or smudges that hide any of the corner marks or the identification number.

If a page is too badly damaged or is illegible, it may be better to just reprint the entire form and transfer the information on the damaged page to the new page; this new page can then replace the bad one.

6.3.2 	Evaluating (or “Verifying”) Forms

Procedure unchanged from Baseline Exam.  Refer to the “Baseline Exam MOP” for further details.

6.3.2.1 	List of scanned images and status of each image

Refer to the “Baseline Exam MOP” for further details.

6.3.2.2 	Making corrections

Procedure unchanged from Baseline Exam.  Refer to the “Baseline Exam MOP” for further details.

6.4 Site Data Management

When you click this icon, you can perform a variety of data management functions including: print forms, update recruitment status (after the initial contact), view a list of participants, view or update participant tracking information, view the status for enrolled participants and look at a summary of forms that have been entered, update data for scanned forms, produce various kinds of reports, send reports to reading centers, transmit data to the Coordinating Center, and run daily data backups.

Each of these functions is listed on the Main Menu that appears when you click the Data Management icon and will be described below.

6.4.1 	Print Forms

Click this button when you want to print out paper Clinic forms.  You can also select the language of the forms that will be printed.  If you change your mind and do not want to print any forms, click on the Close button to exit without printing.  Note that the same selection screen is used for all printing options.

If you select Clinic Forms, you will next be asked to choose between Entire Packet and Specific Participant.  If you choose Entire Packet, you will be printing complete sets of forms for the selected participants.  A screen will come up which lists all of the forms and then asks you to select a Language and a number of forms to print.  Select the appropriate language, and then enter a number between 1 and 5. (Only 5 sets of clinic forms can be printed at one time due to limitations of the printer’s paper tray.)

If you select Clinic Forms and then Specific Participant, you will be asked to enter a Participant ID.  Use this option when one or more of the forms was damaged or lost and needs to be printed again.  Note that the “Number” field is disabled in this option.

# 6.4.2 Participant List

Clicking on the Participant List button on the Main Menu brings up a list of all people who have had a reception form scanned and verified.  Click on “Report” to create a file that can be printed.  Click on the pull down arrow next to Column to sort the list by any one of the columns.  Click on the Participant/Visit Status button or the Tracking Information button to view this information for a participant.  The Participant/Visit Status or Tracking information can also be accessed by clicking the “Tracking” or “Status/Summary” button on the Main Menu.

# 6.4.3 Tracking

Clicking the Tracking button on the Main Menu brings up a screen showing all available contact information for all participants who have had a clinic visit.  This is also the screen that will be used to enter and update the Participant Contact (Tracking) Form information for all participants.

The Visit Summary tab shows the status of the clinic visit, the status of CT and MRI scans, and any notes from the visit.  The Participant Reports tab shows whether the Preliminary and Final participant reports have been mailed, whether the participant reimbursement check has been mailed, and whether the participant had any alerts.

# 6.4.4 	Status/Summary

Click on the Status/Summary button on the Main Menu to display the status of scanned data for a participant.  Note: this screen gives the status of data in your local database, not in the CC database.  If you highlight a form, you can then click on the Update Data button to edit the data from that form. You can click on the Edit Status button to change the status of a form.  Note: the status is set automatically when a form is scanned and verified.  You should use this button only to set the status for forms that will not be completed.  These forms will have a status of “Unknown” or “form missing-no information.”  You can edit the status to one of the following:

2 = not done, physical reason

3 = not done, refused

4 = not done, other reason

5 = not done, cognitive reason

9 = done but form lost

These codes and descriptions are listed when you click on the pull down arrow.

If you have updated the data or status of a form, it will be marked for export to the Coordinating Center automatically.  However, on some occasions you may have to resend data that you have already sent, even though it has not been changed in any way.  In this case, you can click on the Mark for Export button and the data will be sent to the Coordinating Center with the next transmission. After you edit the form status, click Close to close the window and save the information, or Cancel to close without saving.

# 6.4.5 	Reports to Reading Center

Click on this button to send the relevant “Completion Form” to a designated Reading Center or Lab.  Note that this does not transmit the actual scan data file, just the form telling the Reading Center whether or not the procedure was done.

6.4.6 	Coordinating Center Reports

The Coordinating Center will create reports that will be posted on the MESA Web site.  These reports will tell you the status (i.e., whether or not data have been received) of clinic and Reading Center data in the Coordinating Center database.

SECTION 7:	EVENTS SURVEILLANCE

I.	Purpose

MESA’s official, scheduled “Follow-up Phone Calls” serve several purposes:

- To update participants' tracking data, including their address, phone number, and contact information

- To update participants' vital status

- To obtain information regarding participants' general health and health care treatment.

- To obtain detailed information about specific medical conditions diagnosed by a physician that the participants report have had since their last MESA Follow-up phone interview

- To obtain detailed information about any procedures or hospitalizations participants have had since their last MESA Follow-up phone interview

- To schedule the participant for an upcoming Exam Clinic Visit appointment when necessary (Follow-up Call 2, 4, & 6), including MESA CT/MRI appointments when necessary.

7.1 		Follow-up Call

The “Follow-up Call” is a telephone interview scheduled for each participant.  When MESA began, Follow-up Calls were originally scheduled in relation to the dates of the participants’ individual Exam Visits.  Now, however, each Follow-up Call is scheduled in relation to the date of the participant’s previous Follow-up Call and not in relation to the date of the participant’s Exam Visit (even though Follow-up Calls are used to schedule Exam Visits).

If the Exam staff needs to contact a participant (e.g., to schedule an Exam Visit after the Follow-up Call did not succeed in scheduling the Exam Visit), the Exam staff should be aware of the participant’s usual schedule of official Follow-up Calls.  Unscheduled contact is not the same as a scheduled Follow-up Call used to collect data.  If, during an unscheduled contact, the participant tells the Exam staff member about a condition, procedure, or hospital stay, the staff member should alert the Events Coordinator so that full information about the potential event can be collected.

7.2 		Continuing Participant Surveillance

If, during an Exam visit, a participant tells the Exam staff member about a condition, procedure, or hospital stay, the staff member should alert the Events Coordinator so that full information about the potential event can be collected.

At the end of the Exam, each participant is asked to notify the clinic should any changes occur in his/her health, especially involving a hospitalization, nursing home admission, or diagnosis of myocardial infarction (heart attack), angina (chest pain), heart failure (CHF), peripheral vascular disease (PVD), stroke, or transient ischemic attack (TIA).

Clinic staff should inform the participant s/he will be later be contacted for his/her scheduled Follow-up Call, which may include the scheduling of the next Exam Visit.

Let the participant know that we are attempting to keep the record as updated as possible to learn all that we can.

All Follow-up forms are interviewer-administered to MESA participants over the telephone. If the participant prefers to relay this information in person or, for some reason, is unable to complete the interview by phone, a home or clinic visit may be scheduled.

7.1 		Use of Proxy

If a participant is not able to do the interview (e.g., due to a medical problem), a proxy may be used. A “proxy” is a relative or other knowledgeable contact.  If the participant has died, the proxy may complete the questionnaire for the period between the last Follow-up Call and the date of death.

The proxy may or may not be someone previously designated as a contact by the participant. For example, the participant may have designated his/her spouse as a primary contact, but the participant’s son or daughter actually ends up being the person to complete the questionnaire. This is fine, as long as the new person is knowledgeable regarding the participant’s medical condition, procedures of interest, etc.

APPENDIX 1:  Handwriting Examples for Teleform

First, here are two examples of good handwriting for Teleform.  The first set is pretty much “perfect” from a Teleform point of view.  The second set is a bit messier but still acceptable.  Key points are:

- Numbers do not touch any edges of the boxes or go outside the lines

- Numbers are not not written in an ambiguous way.  (Examples of ambiguous numbers are given below.)

Now, here are some examples of writing that Teleform has trouble with:

The Date of Birth was read as:   22 / ~4 /1944.  Problems are:

- 0 goes outside the box

- 4 and 9 can be confused if not written carefully.  If the horizontal “tail” on the number goes past the vertical line (as on the final digit of the year), Teleform will read it as a 4, even if the loop is rounded.  Make sure that 9’s are not confused with 4’s.

Teleform read the Height as 111.7 and the Weight as 117.1, and did not ask for verification of either number.

- 1 and 7 can be confused if not written carefully.  Write the numeral 1 as a straight line with no tail on top or bottome, like:

Teleform read this date as 71 / 20 / 2006, and the only numeral it was unsure of was the 0 in the “Day”.  Problems are:

- 0 in month goes outside the box.  Teleform reads only the part inside the box, which looks most like a 7.

- Second digit in Day has been written over.  It is difficult to tell if it is intended to be a 0 or an 8.  X- out mistakes and write the correction over the box.  Do not write over mistakes as was done here.

- Last two 0’s in the Year look like 6’s, and the last 0 was read as a 6.  Make sure that 0’s are closed at the top.

Teleform read the LAO Aorta number as 89, and the Black Blood Aorta numbers as 7 and 94.  It did not bring any of the numbers up for verification.  Problems are:

- Second digit in the number 88 does not have a loop in the bottom, so it looks like a 9.

- On the Black Blood line, again the first 9 is misread because it goes out of the box, and the last number is ambiguous because it is difficult to tell if the numbers are 4’s or 9’s.

Summary:

  - Make sure that numbers are completely inside the box

  - Form digits carefully so that there are no ambiguities.  In particular, watch for problems with:

1 and 7

4 and 9

6 and 0

8 and 9

3.	See below for optimal printing of letters and numbers:

| A | B | C | D | E | F | G | H | I | J | K | L | # M |  | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 0 |
| --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- | --- |
| N | O | P | Q | R | S | T | U | V | W | X | Y | Z |  |  |  |  |  |  |  |  |  |  |  |